## Test Results
|  |  | Fan Test - V/I for power supplying all 4 fans, RD signal outputs | [7 empty] | Fan PWM Sweep - One-row programmed/read voltage, current, and oscillation frequency | [119 empty] | Heater Test - Results for each heating element and temperature rise after heating time | [7 empty] | HV Test - Results for each configuration. Resistance in units shown, Setup Time k (seconds) in a*e^(-b * t)+c where k = 1/b
Test Name | Date | Time | Supply Voltage | Supply Current | Fan 1 RD | Fan 2 RD | Fan 3 RD | Fan 4 RD | Fan 5 RD | Fan 6 RD | Fan #1 - 12 V : Supply Voltage [V] | Fan #1 - 12 V : Read Voltage [V] | Fan #1 - 12 V : Read Current [A] | Fan #1 - 12 V : Oscillation Frequency [Hz] | Fan #1 - 15 V : Supply Voltage [V] | Fan #1 - 15 V : Read Voltage [V] | Fan #1 - 15 V : Read Current [A] | Fan #1 - 15 V : Oscillation Frequency [Hz] | Fan #1 - 18 V : Supply Voltage [V] | Fan #1 - 18 V : Read Voltage [V] | Fan #1 - 18 V : Read Current [A] | Fan #1 - 18 V : Oscillation Frequency [Hz] | Fan #1 - 21 V : Supply Voltage [V] | Fan #1 - 21 V : Read Voltage [V] | Fan #1 - 21 V : Read Current [A] | Fan #1 - 21 V : Oscillation Frequency [Hz] | Fan #1 - 24 V : Supply Voltage [V] | Fan #1 - 24 V : Read Voltage [V] | Fan #1 - 24 V : Read Current [A] | Fan #1 - 24 V : Oscillation Frequency [Hz] | Fan #2 - 12 V : Supply Voltage [V] | Fan #2 - 12 V : Read Voltage [V] | Fan #2 - 12 V : Read Current [A] | Fan #2 - 12 V : Oscillation Frequency [Hz] | Fan #2 - 15 V : Supply Voltage [V] | Fan #2 - 15 V : Read Voltage [V] | Fan #2 - 15 V : Read Current [A] | Fan #2 - 15 V : Oscillation Frequency [Hz] | Fan #2 - 18 V : Supply Voltage [V] | Fan #2 - 18 V : Read Voltage [V] | Fan #2 - 18 V : Read Current [A] | Fan #2 - 18 V : Oscillation Frequency [Hz] | Fan #2 - 21 V : Supply Voltage [V] | Fan #2 - 21 V : Read Voltage [V] | Fan #2 - 21 V : Read Current [A] | Fan #2 - 21 V : Oscillation Frequency [Hz] | Fan #2 - 24 V : Supply Voltage [V] | Fan #2 - 24 V : Read Voltage [V] | Fan #2 - 24 V : Read Current [A] | Fan #2 - 24 V : Oscillation Frequency [Hz] | Fan #3 - 12 V : Supply Voltage [V] | Fan #3 - 12 V : Read Voltage [V] | Fan #3 - 12 V : Read Current [A] | Fan #3 - 12 V : Oscillation Frequency [Hz] | Fan #3 - 15 V : Supply Voltage [V] | Fan #3 - 15 V : Read Voltage [V] | Fan #3 - 15 V : Read Current [A] | Fan #3 - 15 V : Oscillation Frequency [Hz] | Fan #3 - 18 V : Supply Voltage [V] | Fan #3 - 18 V : Read Voltage [V] | Fan #3 - 18 V : Read Current [A] | Fan #3 - 18 V : Oscillation Frequency [Hz] | Fan #3 - 21 V : Supply Voltage [V] | Fan #3 - 21 V : Read Voltage [V] | Fan #3 - 21 V : Read Current [A] | Fan #3 - 21 V : Oscillation Frequency [Hz] | Fan #3 - 24 V : Supply Voltage [V] | Fan #3 - 24 V : Read Voltage [V] | Fan #3 - 24 V : Read Current [A] | Fan #3 - 24 V : Oscillation Frequency [Hz] | Fan #4 - 12 V : Supply Voltage [V] | Fan #4 - 12 V : Read Voltage [V] | Fan #4 - 12 V : Read Current [A] | Fan #4 - 12 V : Oscillation Frequency [Hz] | Fan #4 - 15 V : Supply Voltage [V] | Fan #4 - 15 V : Read Voltage [V] | Fan #4 - 15 V : Read Current [A] | Fan #4 - 15 V : Oscillation Frequency [Hz] | Fan #4 - 18 V : Supply Voltage [V] | Fan #4 - 18 V : Read Voltage [V] | Fan #4 - 18 V : Read Current [A] | Fan #4 - 18 V : Oscillation Frequency [Hz] | Fan #4 - 21 V : Supply Voltage [V] | Fan #4 - 21 V : Read Voltage [V] | Fan #4 - 21 V : Read Current [A] | Fan #4 - 21 V : Oscillation Frequency [Hz] | Fan #4 - 24 V : Supply Voltage [V] | Fan #4 - 24 V : Read Voltage [V] | Fan #4 - 24 V : Read Current [A] | Fan #4 - 24 V : Oscillation Frequency [Hz] | Fan #5 - 12 V : Supply Voltage [V] | Fan #5 - 12 V : Read Voltage [V] | Fan #5 - 12 V : Read Current [A] | Fan #5 - 12 V : Oscillation Frequency [Hz] | Fan #5 - 15 V : Supply Voltage [V] | Fan #5 - 15 V : Read Voltage [V] | Fan #5 - 15 V : Read Current [A] | Fan #5 - 15 V : Oscillation Frequency [Hz] | Fan #5 - 18 V : Supply Voltage [V] | Fan #5 - 18 V : Read Voltage [V] | Fan #5 - 18 V : Read Current [A] | Fan #5 - 18 V : Oscillation Frequency [Hz] | Fan #5 - 21 V : Supply Voltage [V] | Fan #5 - 21 V : Read Voltage [V] | Fan #5 - 21 V : Read Current [A] | Fan #5 - 21 V : Oscillation Frequency [Hz] | Fan #5 - 24 V : Supply Voltage [V] | Fan #5 - 24 V : Read Voltage [V] | Fan #5 - 24 V : Read Current [A] | Fan #5 - 24 V : Oscillation Frequency [Hz] | Fan #6 - 12 V : Supply Voltage [V] | Fan #6 - 12 V : Read Voltage [V] | Fan #6 - 12 V : Read Current [A] | Fan #6 - 12 V : Oscillation Frequency [Hz] | Fan #6 - 15 V : Supply Voltage [V] | Fan #6 - 15 V : Read Voltage [V] | Fan #6 - 15 V : Read Current [A] | Fan #6 - 15 V : Oscillation Frequency [Hz] | Fan #6 - 18 V : Supply Voltage [V] | Fan #6 - 18 V : Read Voltage [V] | Fan #6 - 18 V : Read Current [A] | Fan #6 - 18 V : Oscillation Frequency [Hz] | Fan #6 - 21 V : Supply Voltage [V] | Fan #6 - 21 V : Read Voltage [V] | Fan #6 - 21 V : Read Current [A] | Fan #6 - 21 V : Oscillation Frequency [Hz] | Fan #6 - 24 V : Supply Voltage [V] | Fan #6 - 24 V : Read Voltage [V] | Fan #6 - 24 V : Read Current [A] | Fan #6 - 24 V : Oscillation Frequency [Hz] | TC1_Resistance | TC2_Resistance | TC3_Resistance | TC4_Resistance | TC1_Temp_Rise | TC2_Temp_Rise | TC3_Temp_Rise | TC4_Temp_Rise | Ch0+ Open Res | Ch0+ Open Fit On | Ch0+ Open Fit Off | Ch0+ 10k Res | Ch0+ 10k Fit On | Ch0+ 10k Fit Off | Ch0- Open Res | Ch0- Open Fit On | Ch0- Open Fit Off | Ch0- 10k Res | Ch0- 10k Fit On | Ch0- 10k Fit Off | Ch1+ Open Res | Ch1+ Open Fit On | Ch1+ Open Fit Off | Ch1+ 10k Res | Ch1+ 10k Fit On | Ch1+ 10k Fit Off | Ch1- Open Res | Ch1- Open Fit On | Ch1- Open Fit Off | Ch1- 10k Res | Ch1- 10k Fit On | Ch1- 10k Fit Off | Ch2+ Open Res | Ch2+ Open Fit On | Ch2+ Open Fit Off | Ch2+ 10k Res | Ch2+ 10k Fit On | Ch2+ 10k Fit Off | Ch2- Open Res | Ch2- Open Fit On | Ch2- Open Fit Off | Ch2- 10k Res | Ch2- 10k Fit On | Ch2- 10k Fit Off | Ch3+ Open Res | Ch3+ Open Fit On | Ch3+ Open Fit Off | Ch3+ 10k Res | Ch3+ 10k Fit On | Ch3+ 10k Fit Off | Ch3- Open Res | Ch3- Open Fit On | Ch3- Open Fit Off | Ch3- 10k Res | Ch3- 10k Fit On | Ch3- 10k Fit Off | Ch4+ Open Res | Ch4+ Open Fit On | Ch4+ Open Fit Off | Ch4+ 10k Res | Ch4+ 10k Fit On | Ch4+ 10k Fit Off | Ch4- Open Res | Ch4- Open Fit On | Ch4- Open Fit Off | Ch4- 10k Res | Ch4- 10k Fit On | Ch4- 10k Fit Off | Ch5+ Open Res | Ch5+ Open Fit On | Ch5+ Open Fit Off | Ch5+ 10k Res | Ch5+ 10k Fit On | Ch5+ 10k Fit Off | Ch5- Open Res | Ch5- Open Fit On | Ch5- Open Fit Off | Ch5- 10k Res | Ch5- 10k Fit On | Ch5- 10k Fit Off | Ch6+ Open Res | Ch6+ Open Fit On | Ch6+ Open Fit Off | Ch6+ 10k Res | Ch6+ 10k Fit On | Ch6+ 10k Fit Off | Ch6- Open Res | Ch6- Open Fit On | Ch6- Open Fit Off | Ch6- 10k Res | Ch6- 10k Fit On | Ch6- 10k Fit Off | Ch7+ Open Res | Ch7+ Open Fit On | Ch7+ Open Fit Off | Ch7+ 10k Res | Ch7+ 10k Fit On | Ch7+ 10k Fit Off | Ch7- Open Res | Ch7- Open Fit On | Ch7- Open Fit Off | Ch7- 10k Res | Ch7- 10k Fit On | Ch7- 10k Fit Off
slow_ctl_crate_with_wibs | 05/18/2026 | 10:23:02 AM | 23.9904 | 1.1549 | 0.18 | 0.19 | 0.19 | 0.18 | 0.18 | 0.18 | [120 empty] | 18.17 | 18.84 | 18.3 | 17.86 | 1.87 | 1.93 | 2.01 | 4.2 | 5998.6Mohm | 0.02 | 0.02 | 59.86kohm | 0.04 | 0.01 | 7497.15Mohm | 0.02 | 0.02 | 57.83kohm | 0.01 | 1 | 6001.56Mohm | 0.02 | 0.02 | 58.53kohm | 0 | 9.83 | 5999.44Mohm | 0.02 | 0.02 | 57.52kohm | 0.02 | 1 | 5998.24Mohm | 0.02 | 0.02 | 58.08kohm | 0.01 | 0.02 | 7891.79Mohm | 0.02 | 0.02 | 58.98kohm | 0.02 | 0.03 | 5770.08Mohm | 0.02 | 0.02 | 58.33kohm | 0.03 | 0.03 | 7500.65Mohm | 0.02 | 0.02 | 59.4kohm | 0.04 | 37.02 | 5768.12Mohm | 0.02 | 0.03 | 109.1kohm | 0.12 | 0.02 | 5997.72Mohm | 0.01 | 0.03 | 110.09kohm | 0.01 | 1 | 4412.44Mohm | 0.02 | 0.03 | 105.2kohm | 0.03 | 1 | 5001.03Mohm | 0.01 | 0.03 | 111.0kohm | 0.02 | 0.04 | 7141.19Mohm | 0.02 | 0.03 | 110.22kohm | 0.01 | 0.01 | 4410.26Mohm | 0.01 | 0.03 | 106.66kohm | 0.09 | 1 | 7143.81Mohm | 0.02 | 0.03 | 113.75kohm | 0 | 0.04 | 6523.09Mohm | 0.02 | 0.03 | 108.8kohm | 0.05 | 1
slow_ctl_crate_with_wibs | 05/18/2026 | 05:19:05 PM | 23.9881 | 1.1098 | 0.18 | 0.19 | 0.19 | 0.18 | 0.18 | 0.18 | [120 empty] | 18.2 | 18.87 | 18.32 | 17.88 | 1.86 | 1.91 | 2 | 4.18 | 7140.62Mohm | 0.02 | 0.02 | 59.51kohm | 0.06 | 0.02 | 8330.83Mohm | 0.02 | 0.02 | 57.95kohm | 0.04 | 1 | 7144.76Mohm | 0.02 | 0.02 | 58.73kohm | 0.04 | 2.03 | 7142.48Mohm | 0.02 | 0.02 | 57.8kohm | 0 | 1 | 7498.2Mohm | 0.02 | 0.02 | 58.33kohm | 0.01 | 0.02 | 11533.92Mohm | 0.02 | 0.02 | 58.44kohm | 0.02 | 0.03 | 7501.2Mohm | 0.02 | 0.02 | 59.08kohm | 0.01 | 0.02 | 9375.75Mohm | 0.02 | 0.02 | 59.19kohm | 0.02 | 0.03 | 9373.81Mohm | 0.02 | 0.03 | 109.6kohm | 0.03 | 0.02 | 8820.71Mohm | 0.02 | 0.03 | 109.38kohm | 0.04 | 1 | 7144.33Mohm | 0.02 | 0.03 | 104.89kohm | 0.03 | 0.13 | 7502.1Mohm | 0.02 | 0.03 | 110.41kohm | 0.02 | 0.04 | 9998.33Mohm | 0.02 | 0.03 | 109.7kohm | 0.04 | 0.02 | 5355.14Mohm | 0.01 | 0.03 | 105.63kohm | 0.02 | 1 | 10716.29Mohm | 0.02 | 0.03 | 113.88kohm | 0.02 | 0.03 | 8825.29Mohm | 0.01 | 0.03 | 108.82kohm | 0.04 | 0.12
slow_ctl_crate_with_delta | 05/19/2026 | 03:45:37 PM | 23.9894 | 1.1432 | 2.52 | 2.51 | 2.51 | 2.5 | 2.48 | 2.51 | [120 empty] | 18.2 | 18.93 | 18.34 | 17.91 | 1.83 | 1.85 | 1.98 | 4.19 | 7892.79Mohm | 0.02 | 0.02 | 58.78kohm | 0 | 0.03 | 9371.56Mohm | 0.02 | 0.02 | 57.47kohm | 0.01 | 1 | 8336.06Mohm | 0.02 | 0.02 | 58.15kohm | 0.03 | 0.04 | 7499.65Mohm | 0.02 | 0.02 | 57.28kohm | 0.03 | 0.04 | 8331.0Mohm | 0.02 | 0.02 | 58.46kohm | 0.03 | 0.02 | 14994.5Mohm | 0.02 | 0.02 | 58.78kohm | 0.03 | 0.02 | 8334.56Mohm | 0.02 | 0.02 | 58.94kohm | 0.01 | 0.02 | 11538.77Mohm | 0.02 | 0.02 | 59.37kohm | 0.02 | 0 | 10712.57Mohm | 0.02 | 0.03 | 109.01kohm | 0.01 | 0 | 9996.13Mohm | 0.01 | 0.03 | 108.69kohm | 0.04 | 1 | 7144.19Mohm | 0.02 | 0.03 | 104.26kohm | 0.01 | 1 | 7896.53Mohm | 0.02 | 0.03 | 110.74kohm | 0.04 | 0.04 | 13633.64Mohm | 0.02 | 0.03 | 109.97kohm | 0.02 | 0.03 | 5997.92Mohm | 0.01 | 0.03 | 106.78kohm | 0.02 | 1 | 13638.27Mohm | 0.02 | 0.03 | 114.09kohm | 0.02 | 0.16 | 13639.09Mohm | 0.01 | 0.03 | 108.79kohm | 0.02 | 1
disconnected_test | 06/02/2026 | 01:54:08 PM | 23.9902 | 0 | 0 | 0 | 0 | 0 | 0 | 0
disconnected_test | 06/02/2026 | 01:55:19 PM | 23.9902 | 0 | 0 | 0 | 0 | 0 | 0 | 0
fan_connected_test | 06/02/2026 | 02:10:32 PM | 23.9908 | 1.3161 | 303.52 | 311.12 | 320.11 | 324.47 | 298.21 | 315.19
fan_connected_test | 06/02/2026 | 02:15:22 PM | 23.9905 | 1.3142 | 304.02 | 312.63 | 320.49 | 325.14 | 298.05 | 315.1
fan_connected_test | 06/02/2026 | 02:16:56 PM
fan_connected_test | 06/02/2026 | 02:17:35 PM
fan_connected_test | 06/02/2026 | 02:18:11 PM | 23.9904 | 1.3068 | 305.06 | 313.29 | 320.21 | 325.4 | 298.62 | 316.36
old_fxns_test | 06/02/2026 | 02:21:57 PM | 23.9904 | 1.3093 | 2.58 | 2.57 | 2.59 | 2.4 | 2.57 | 2.57
old_fxns_test_disconnected | 06/03/2026 | 12:57:38 PM | 23.9904 | 0 | 5 | 5 | 5 | 5 | 5 | 5
skip_fan_test | 06/03/2026 | 12:59:38 PM
skip_fan_test_rtd_delay_10 | 06/03/2026 | 01:01:04 PM | [128 empty] | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 0 | 0 | 0 | 0
skip_fan_test_rtd_delay_6 | 06/03/2026 | 01:04:58 PM | [128 empty] | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 0 | 0 | 0 | 0
skip_fan_test_rtd_delay_1 | 06/03/2026 | 01:07:25 PM
skip_fan_test_rtd_delay_5 | 06/03/2026 | 01:08:00 PM | [128 empty] | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 0 | 0 | 0 | 0
rtd_delay_5_uncomment_fan_test | 06/03/2026 | 01:09:39 PM | 23.9904 | 0 | 5 | 5 | 5 | 5 | 5 | 5 | [120 empty] | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 0 | 0 | 0 | 0
test_connected_fandelay_4 | 06/04/2026 | 02:57:36 PM | 23.9903 | 1.3096 | 307.87 | 324.56 | 308.97 | 319.43 | 299.96 | 315.27 | [120 empty] | 18.14 | 18.89 | 18.34 | 17.85 | 1.71 | 1.76 | 1.87 | 3.68
test_integrated_code | 06/05/2026 | 03:22:42 PM
test_integrated_code | 06/05/2026 | 03:23:25 PM | 23.9904 | 1.3157 | 306.57 | 315.59 | 312.07 | 318.02 | 305.3 | 313.19 | [120 empty] | 18.14 | 18.86 | 18.32 | 17.85 | 1.73 | 1.79 | 1.89 | 3.82
quick_test | 06/09/2026 | 01:52:50 PM | 23.9906 | 1.3187 | 305.05 | 314.33 | 312.46 | 316.86 | 303.66 | 312.83 | [120 empty] | 18.15 | 18.9 | 18.34 | 17.85 | 1.73 | 1.78 | 1.88 | 3.79
001_Jillian | 06/09/2026 | 02:25:28 PM | 23.9907 | 1.3156 | 305.53 | 315.51 | 312.94 | 318.5 | 305.25 | 312.44 | [120 empty] | 18.16 | 18.92 | 18.34 | 17.86 | 1.73 | 1.78 | 1.89 | 3.81 | 10711.64Mohm | 0.02 | 0.03 | 111.73kohm | 0.01 | 0.02 | 12495.75Mohm | 0.02 | 0.03 | 108.74kohm | 0 | 1 | 8335.94Mohm | 0.02 | 0.03 | 110.65kohm | 0.04 | 0.04 | 7895.32Mohm | 0.02 | 0.03 | 108.16kohm | 0.03 | 0.04 | 13632.55Mohm | 0.02 | 0.03 | 110.31kohm | 0.01 | 0.02 | 24990.17Mohm | 0.02 | 0.03 | 110.04kohm | 0 | 0.03 | 10716.5Mohm | 0.02 | 0.03 | 110.9kohm | 0.03 | 0.01 | 18752.5Mohm | 0.02 | 0.03 | 112.41kohm | 0 | 0 | 14998.1Mohm | 0.02 | 0.02 | 58.12kohm | 0.06 | 0 | 11534.38Mohm | 0.02 | 0.02 | 57.84kohm | 0.04 | 1 | 10716.71Mohm | 0.02 | 0.02 | 55.7kohm | 0.05 | 0.43 | 15005.1Mohm | 0.02 | 0.02 | 58.25kohm | 0.02 | 0.04 | 10712.36Mohm | 0.02 | 0.02 | 58.48kohm | 0.06 | 0.01 | 5767.04Mohm | 0.02 | 0.02 | 55.31kohm | 0.03 | 1 | 9376.63Mohm | 0.02 | 0.02 | 60.52kohm | 0 | 0.05 | 7896.95Mohm | 0.02 | 0.02 | 57.31kohm | 0.01 | 1
001_Jillian_shutoff | 06/09/2026 | 03:01:45 PM | 23.9905 | 1.3136 | 305.57 | 316.08 | 313.04 | 318.99 | 305.24 | 313.04 | [120 empty] | 18.16 | 18.91 | 18.33 | 17.85 | 1.72 | 1.77 | 1.88 | 3.8
001_Jillian_abort | 06/09/2026 | 03:19:29 PM
001_amshi_hibaysetup_connection_test | 06/11/2026 | 02:58:41 PM | 23.9902 | 1.307 | 316.65 | 310.44 | 315.16 | 317.94 | 305.76 | 315.71 | [120 empty] | 18.14 | 18.85 | 18.29 | 17.84 | 1.72 | 1.78 | 1.88 | 3.8 | 8821.41Mohm | 0.01 | 0.03 | 110.46kohm | 0.01 | 17.45 | 13631.82Mohm | 0.02 | 0.03 | 107.88kohm | 0.02 | 1 | 7145.1Mohm | 0.01 | 0.03 | 109.87kohm | 0.02 | 0.04 | 7500.1Mohm | 0.02 | 0.03 | 109.36kohm | 0.02 | 1 | 13632.64Mohm | 0.02 | 0.03 | 110.35kohm | 0.02 | 0.03 | 0.0Mohm | 0.02 | 0.03 | 109.92kohm | 0.03 | 0.02 | 11540.08Mohm | 0.02 | 0.03 | 110.7kohm | 0.03 | 0.03 | 49999.67Mohm | 0.02 | 0.03 | 112.44kohm | 0.02 | 0 | 24997.67Mohm | 0.02 | 0.02 | 58.11kohm | 0.28 | 1 | 16661.56Mohm | 0.02 | 0.02 | 57.48kohm | 0.01 | 0.04 | 12503.25Mohm | 0.02 | 0.02 | 55.04kohm | 0 | 1 | 18754.88Mohm | 0.02 | 0.02 | 58.19kohm | 0.06 | 0 | 13633.27Mohm | 0.02 | 0.02 | 58.77kohm | 0.01 | 0.02 | 6815.59Mohm | 0.02 | 0.02 | 55.96kohm | 0.05 | 1 | 10716.0Mohm | 0.02 | 0.02 | 60.32kohm | 0.02 | 0.04 | 16669.89Mohm | 0.02 | 0.02 | 57.6kohm | 0.04 | 1
001_amshi_fan_pwm_formatfixcheck | 06/23/2026 | 10:48:34 AM | 23.9907 | 1.3083 | 324.5684 | 300.4273 | 306.6657 | 320.3158 | 306.5468 | 320.1943 | 12 | 11.9888 | 0.704 | 191.6449 | 15 | 14.9889 | 0.8724 | 233.1569 | 18 | 17.989 | 1.0155 | 267.9099 | 21 | 20.9884 | 1.1577 | 297.8711 | 24 | 23.9902 | 1.3054 | 324.9136 | 12 | 11.9888 | 0.704 | 177.4789 | 15 | 14.9889 | 0.8724 | 214.746 | 18 | 17.989 | 1.0155 | 246.5418 | 21 | 20.9884 | 1.1577 | 274.8851 | 24 | 23.9902 | 1.3054 | 300.3791 | 12 | 11.9888 | 0.704 | 182.05 | 15 | 14.9889 | 0.8724 | 219.6499 | 18 | 17.989 | 1.0155 | 252.0617 | 21 | 20.9884 | 1.1577 | 280.5301 | 24 | 23.9902 | 1.3054 | 306.5613 | 12 | 11.9888 | 0.704 | 190.9508 | 15 | 14.9889 | 0.8724 | 229.7957 | 18 | 17.989 | 1.0155 | 262.8376 | 21 | 20.9884 | 1.1577 | 293.0157 | 24 | 23.9902 | 1.3054 | 319.6799 | 12 | 11.9888 | 0.704 | 181.0261 | 15 | 14.9889 | 0.8724 | 218.695 | 18 | 17.989 | 1.0155 | 251.187 | 21 | 20.9884 | 1.1577 | 279.7424 | 24 | 23.9902 | 1.3054 | 306.0211 | 12 | 11.9888 | 0.704 | 188.6045 | 15 | 14.9889 | 0.8724 | 228.3753 | 18 | 17.989 | 1.0155 | 262.5999 | 21 | 20.9884 | 1.1577 | 292.5263 | 24 | 23.9902 | 1.3054 | 319.5658 | 18.7398 | 18.2223 | 18.2858 | 17.8551 | 1.7604 | 1.7632 | 1.9096 | 3.8033 | 4345.86943Mohm | 36.1382 | 31.8588 | 112.95884kohm | 65.2912 | 40.5964 | 5879.4117Mohm | 54.2251 | 31.7178 | 109.15453kohm | 30.8901 | 1 | 3573.53579Mohm | 28.2003 | 32.64 | 111.11713kohm | 38.1048 | 24.3663 | 3846.00023Mohm | 45.2196 | 32.1666 | 109.73028kohm | 110.941 | 23.44 | 7689.61585Mohm | 40.816 | 33.1039 | 110.72046kohm | 239.4559 | 36.9316 | 19989.79951Mohm | 55.5837 | 32.5447 | 111.2602kohm | 26.9857 | 46.634 | 7146.42854Mohm | 51.0461 | 33.4279 | 111.53677kohm | 34.4142 | 32.791 | 12497.12552Mohm | 51.4967 | 32.686 | 111.57047kohm | 48.8643 | 27.7444 | 9090.18171Mohm | 44.9352 | 32.222 | 110.55499kohm | 52.1814 | 219169.0602 | 7140.28579Mohm | 47.6268 | 32.1919 | 109.55582kohm | 58.3642 | 1 | 11115.33295Mohm | 45.9502 | 32.8412 | 104.91354kohm | 294.0068 | 4.8318 | 12505.00058Mohm | 45.018 | 33.1975 | 111.53038kohm | 17.851 | 23.9038 | 5261.21062Mohm | 2.1841 | 33.083 | 110.82617kohm | 97.6072 | 45.5154 | 4543.50006Mohm | 77.8392 | 32.7362 | 106.96586kohm | 0 | 1 | 5557.83296Mohm | 0 | 33.3914 | 113.88515kohm | 25.2346 | 22.652 | 5883.11761Mohm | 0 | 32.9854 | 108.5305kohm | 0 | 1
001_amshi_fan_pwm_passcheck | 06/23/2026 | 11:23:40 AM | 23.9903 | 1.305 | 325.059 | 300.8398 | 306.6955 | 320.6838 | 306.381 | 320.2719 | 12 | 11.9881 | 0.6993 | 192.5002 | 15 | 14.9882 | 0.8692 | 233.8996 | 18 | 17.9884 | 1.0141 | 268.0378 | 21 | 20.9879 | 1.1567 | 297.6273 | 24 | 23.9899 | 1.3051 | 324.8104 | 12 | 11.9881 | 0.6993 | 177.3824 | 15 | 14.9882 | 0.8692 | 214.7791 | 18 | 17.9884 | 1.0141 | 246.6035 | 21 | 20.9879 | 1.1567 | 274.917 | 24 | 23.9899 | 1.3051 | 300.5324 | 12 | 11.9881 | 0.6993 | 181.8928 | 15 | 14.9882 | 0.8692 | 219.8227 | 18 | 17.9884 | 1.0141 | 252.1832 | 21 | 20.9879 | 1.1567 | 280.5498 | 24 | 23.9899 | 1.3051 | 306.8125 | 12 | 11.9881 | 0.6993 | 190.5986 | 15 | 14.9882 | 0.8692 | 231.0797 | 18 | 17.9884 | 1.0141 | 263.0341 | 21 | 20.9879 | 1.1567 | 292.8008 | 24 | 23.9899 | 1.3051 | 319.9689 | 12 | 11.9881 | 0.6993 | 181.3297 | 15 | 14.9882 | 0.8692 | 219.4751 | 18 | 17.9884 | 1.0141 | 251.4242 | 21 | 20.9879 | 1.1567 | 280.136 | 24 | 23.9899 | 1.3051 | 306.4175 | 12 | 11.9881 | 0.6993 | 188.5898 | 15 | 14.9882 | 0.8692 | 228.2879 | 18 | 17.9884 | 1.0141 | 262.5634 | 21 | 20.9879 | 1.1567 | 292.1123 | 24 | 23.9899 | 1.3051 | 319.6138 | 18.7422 | 18.2251 | 18.2889 | 17.8575 | 1.7546 | 1.7577 | 1.9042 | 3.7884 | 4997.79991Mohm | 64.8006 | 31.9024 | 111.95772kohm | 37.8941 | 45.8579 | 5260.57906Mohm | 55.9547 | 31.5469 | 108.28981kohm | 19.121 | 1 | 3848.46159Mohm | 56.8615 | 32.6778 | 111.00276kohm | 0 | 0 | 3448.03456Mohm | 55.1157 | 31.9083 | 109.20446kohm | 212.918 | 22.5931 | 7140.28579Mohm | 50.3759 | 33.0583 | 110.75755kohm | 23.0569 | 42.9303 | 24988.25129Mohm | 68.9807 | 32.311 | 111.17321kohm | 30.4501 | 46.7747 | 6669.86627Mohm | 45.6368 | 33.3863 | 111.32395kohm | 22.6049 | 48.0048 | 12497.62525Mohm | 60.9602 | 32.6142 | 111.28655kohm | 29.9381 | 23.6708 | 9090.63615Mohm | 49.2028 | 32.1986 | 109.12125kohm | 21.4071 | 0 | 7140.07129Mohm | 52.808 | 32.1571 | 109.60489kohm | 0 | 1 | 10003.9Mohm | 46.5115 | 32.8323 | 104.45006kohm | 39.2663 | 1 | 12504.75034Mohm | 49.4846 | 33.1217 | 110.26596kohm | 40.0274 | 25.3855 | 6664.33323Mohm | 12.2414 | 33.0675 | 111.66834kohm | 1722.8257 | 31.1124 | 3569.60709Mohm | 72.2218 | 32.5765 | 107.11898kohm | 60.4151 | 1 | 8336.83338Mohm | 72.2623 | 33.4007 | 113.41483kohm | 36.4677 | 26.8707 | 6250.93779Mohm | 70.0733 | 32.9093 | 109.79245kohm | 23.3559 | 1
001_amshi_test2 | 06/24/2026 | 09:15:17 AM | 23.9917 | 1.3123 | 324.0511 | 300.0267 | 305.4491 | 320.0081 | 305.9016 | 319.1463 | 12 | 11.9907 | 0.7095 | 191.6832 | 15 | 14.9902 | 0.8734 | 232.7546 | 18 | 17.99 | 1.0181 | 267.6837 | 21 | 20.9891 | 1.1591 | 297.547 | 24 | 23.9906 | 1.3069 | 324.5204 | 12 | 11.9907 | 0.7095 | 177.267 | 15 | 14.9902 | 0.8734 | 214.5377 | 18 | 17.99 | 1.0181 | 246.4475 | 21 | 20.9891 | 1.1591 | 274.6513 | 24 | 23.9906 | 1.3069 | 300.4875 | 12 | 11.9907 | 0.7095 | 181.9633 | 15 | 14.9902 | 0.8734 | 219.7628 | 18 | 17.99 | 1.0181 | 251.5122 | 21 | 20.9891 | 1.1591 | 279.908 | 24 | 23.9906 | 1.3069 | 306.1561 | 12 | 11.9907 | 0.7095 | 189.9039 | 15 | 14.9902 | 0.8734 | 229.8417 | 18 | 17.99 | 1.0181 | 263.0065 | 21 | 20.9891 | 1.1591 | 293.0823 | 24 | 23.9906 | 1.3069 | 320.1973 | 12 | 11.9907 | 0.7095 | 181.0937 | 15 | 14.9902 | 0.8734 | 219.1103 | 18 | 17.99 | 1.0181 | 251.0398 | 21 | 20.9891 | 1.1591 | 279.656 | 24 | 23.9906 | 1.3069 | 305.9476 | 12 | 11.9907 | 0.7095 | 186.819 | 15 | 14.9902 | 0.8734 | 227.4993 | 18 | 17.99 | 1.0181 | 261.712 | 21 | 20.9891 | 1.1591 | 291.7785 | 24 | 23.9906 | 1.3069 | 319.1856 | 18.732 | 18.2238 | 18.2895 | 17.8439 | 1.7558 | 1.7569 | 1.9082 | 3.8015 | 5552.94439Mohm | 53.4428 | 32.0548 | 112.14588kohm | 32.5388 | 36.3443 | 6662.99982Mohm | 44.0278 | 31.9126 | 111.47559kohm | 17.4043 | 1 | 4350.13048Mohm | 47.9742 | 32.8959 | 111.56352kohm | 54.2398 | 22.9062 | 4166.20849Mohm | 46.5544 | 32.3615 | 108.33107kohm | 99.3009 | 1 | 9087.45455Mohm | 59.0164 | 33.1729 | 110.57584kohm | 160.7979 | 87.0022 | 33315.66643Mohm | 63.4838 | 32.7407 | 111.15455kohm | 31.9024 | 32.5929 | 7695.46161Mohm | 51.4251 | 33.6567 | 110.47085kohm | 143.3972 | 55.8661 | 12496.00019Mohm | 46.4646 | 32.8828 | 112.38407kohm | 25.3874 | 24.0174 | 11109.55495Mohm | 63.3245 | 32.3998 | 109.86986kohm | 0 | 1 | 7689.07734Mohm | 46.128 | 32.4192 | 108.86006kohm | 19.9402 | 1 | 11113.77723Mohm | 52.0497 | 32.9925 | 104.8596kohm | 0 | 9.8341 | 12504.75034Mohm | 43.5891 | 33.3549 | 110.22942kohm | 0.67306 | 23.5868 | 7139.92873Mohm | 389.5137 | 33.2752 | 111.70777kohm | 40.5696 | 64.0241 | 3998.28003Mohm | 0 | 32.8581 | 108.59627kohm | 5.4193 | 1 | 7145.35694Mohm | 176.7383 | 33.7127 | 113.70893kohm | 78.0184 | 0.03051 | 7693.38501Mohm | 389.4968 | 33.2393 | 107.70483kohm | 63.5541 | 1
002_1500V_test | 06/25/2026 | 02:50:17 PM | 23.99 | 1.3173 | 325.8934 | 300.8329 | 306.9714 | 320.564 | 307.0276 | 316.4541 | 12 | 11.988 | 0.7049 | 192.91 | 15 | 14.988 | 0.8708 | 234.3823 | 18 | 17.9881 | 1.0169 | 268.7581 | 21 | 20.9876 | 1.1613 | 298.7266 | 24 | 23.9897 | 1.3122 | 326.1339 | 12 | 11.988 | 0.7049 | 177.8202 | 15 | 14.988 | 0.8708 | 214.4499 | 18 | 17.9881 | 1.0169 | 247.1889 | 21 | 20.9876 | 1.1613 | 274.8934 | 24 | 23.9897 | 1.3122 | 301.8663 | 12 | 11.988 | 0.7049 | 182.883 | 15 | 14.988 | 0.8708 | 220.8109 | 18 | 17.9881 | 1.0169 | 252.7236 | 21 | 20.9876 | 1.1613 | 280.8425 | 24 | 23.9897 | 1.3122 | 307.3028 | 12 | 11.988 | 0.7049 | 189.4275 | 15 | 14.988 | 0.8708 | 231.0759 | 18 | 17.9881 | 1.0169 | 268.7501 | 21 | 20.9876 | 1.1613 | 294.5802 | 24 | 23.9897 | 1.3122 | 321.3012 | 12 | 11.988 | 0.7049 | 182.3899 | 15 | 14.988 | 0.8708 | 220.2201 | 18 | 17.9881 | 1.0169 | 252.4209 | 21 | 20.9876 | 1.1613 | 280.7701 | 24 | 23.9897 | 1.3122 | 307.0117 | 12 | 11.988 | 0.7049 | 188.7181 | 15 | 14.988 | 0.8708 | 227.1333 | 18 | 17.9881 | 1.0169 | 260.2606 | 21 | 20.9876 | 1.1613 | 289.293 | 24 | 23.9897 | 1.3122 | 316.58 | 18.1375 | 18.8293 | 17.859 | 18.2816 | 1.7124 | 1.7878 | 3.7918 | 1.9609 | 8821.11735Mohm | 56.4258 | 31.8514 | 113.3992kohm | 39.9468 | 98.7525 | 11534.53881Mohm | 52.9346 | 31.7486 | 108.51595kohm | 27.9111 | 1 | 7145.23832Mohm | 57.7048 | 32.8293 | 110.86403kohm | 29.4988 | 270744.5639 | 7143.38111Mohm | 47.193 | 32.4278 | 109.08106kohm | 0 | 24.2706 | 11535.38485Mohm | 61.6762 | 32.5906 | 110.54603kohm | 56.7262 | 58.652 | 49978.33364Mohm | 59.2972 | 32.2341 | 111.60091kohm | 0 | 44.1634 | 9376.68783Mohm | 62.5014 | 32.9043 | 110.21594kohm | 10.2977 | 42.309 | 25001.00154Mohm | 59.2535 | 32.375 | 112.12105kohm | 0 | 23.6556 | 37494.25133Mohm | 49.6144 | 32.8049 | 109.60652kohm | 26.1801 | 0 | 21421.42848Mohm | 54.3392 | 33.233 | 109.74466kohm | 0 | 1 | 16670.4434Mohm | 50.5242 | 31.9529 | 105.0826kohm | 13.3748 | 1 | 21436.28618Mohm | 53.7722 | 32.648 | 110.80004kohm | 43.5651 | 27.47 | 18747.50103Mohm | 48.3326 | 32.5327 | 110.1065kohm | 71.733 | 41.1811 | 7140.00034Mohm | 55.4177 | 32.4198 | 107.30518kohm | 0 | 1 | 15002.69997Mohm | 50.4035 | 32.873 | 114.63322kohm | 26.9919 | 23.6439 | 11540.61603Mohm | 55.0467 | 32.7716 | 108.05543kohm | 0 | 24.2554
002_1800V_test | 06/25/2026 | 03:57:41 PM | 23.99 | 1.3133 | 326.4758 | 301.8485 | 307.5557 | 321.336 | 307.1688 | 316.8912 | 12 | 11.9879 | 0.7024 | 193.16 | 15 | 14.9881 | 0.8714 | 234.5726 | 18 | 17.9884 | 1.0169 | 268.9884 | 21 | 20.9877 | 1.1604 | 298.9204 | 24 | 23.9896 | 1.3115 | 326.3326 | 12 | 11.9879 | 0.7024 | 178.031 | 15 | 14.9881 | 0.8714 | 214.6068 | 18 | 17.9884 | 1.0169 | 247.1799 | 21 | 20.9877 | 1.1604 | 276.0612 | 24 | 23.9896 | 1.3115 | 301.7464 | 12 | 11.9879 | 0.7024 | 182.8909 | 15 | 14.9881 | 0.8714 | 220.7433 | 18 | 17.9884 | 1.0169 | 252.8414 | 21 | 20.9877 | 1.1604 | 281.3203 | 24 | 23.9896 | 1.3115 | 307.4979 | 12 | 11.9879 | 0.7024 | 189.4027 | 15 | 14.9881 | 0.8714 | 231.9409 | 18 | 17.9884 | 1.0169 | 264.507 | 21 | 20.9877 | 1.1604 | 293.8568 | 24 | 23.9896 | 1.3115 | 321.3506 | 12 | 11.9879 | 0.7024 | 182.4197 | 15 | 14.9881 | 0.8714 | 220.3056 | 18 | 17.9884 | 1.0169 | 252.1566 | 21 | 20.9877 | 1.1604 | 280.7151 | 24 | 23.9896 | 1.3115 | 307.2071 | 12 | 11.9879 | 0.7024 | 188.8111 | 15 | 14.9881 | 0.8714 | 227.4124 | 18 | 17.9884 | 1.0169 | 260.3203 | 21 | 20.9877 | 1.1604 | 289.5098 | 24 | 23.9896 | 1.3115 | 316.8349 | 18.1435 | 18.8363 | 17.8657 | 18.2875 | 1.7174 | 1.7891 | 3.795 | 1.9655 | 11247.43754Mohm | 62.3049 | 31.5441 | 111.91256kohm | 36.3434 | 27.0611 | 13841.07717Mohm | 59.5 | 31.4574 | 108.91144kohm | 84.2062 | 1 | 9001.5996Mohm | 64.6849 | 32.5602 | 111.29163kohm | 35.298 | 24.2293 | 8572.23857Mohm | 58.8514 | 32.1857 | 108.44941kohm | 20.5017 | 25.3423 | 16359.27299Mohm | 71.2718 | 32.3521 | 110.26085kohm | 19.4668 | 45.9733 | 0.0Mohm | 66.8567 | 31.9883 | 111.32223kohm | 29.0349 | 28.0576 | 12000.13299Mohm | 62.6196 | 32.6587 | 110.49119kohm | 33.2435 | 127.5567 | 60009.00199Mohm | 62.5279 | 32.0077 | 111.62538kohm | 28.3641 | 0.03918 | 89983.00372Mohm | 57.5382 | 32.6256 | 110.58198kohm | 24.0638 | 85.6207 | 29990.83319Mohm | 56.6526 | 32.605 | 110.58486kohm | 20.0999 | 1 | 20001.33248Mohm | 54.9166 | 31.8861 | 105.26593kohm | 0 | 1 | 30006.66774Mohm | 54.336 | 32.1338 | 110.42978kohm | 24.7188 | 29.3097 | 22496.62526Mohm | 55.3136 | 32.2137 | 110.35437kohm | 19.7526 | 30.2548 | 9470.15778Mohm | 57.6775 | 31.9775 | 106.50559kohm | 51.5492 | 1 | 20002.10967Mohm | 50.9869 | 32.4802 | 113.47175kohm | 0 | 23.8481 | 15001.58318Mohm | 54.7789 | 32.2561 | 108.74858kohm | 38.4694 | 0.10154
002_2000V_test | 06/25/2026 | 04:30:38 PM | 23.9901 | 1.3129 | 326.7618 | 301.8547 | 307.3653 | 321.0832 | 307.4106 | 317.1029 | 12 | 11.988 | 0.7008 | 193.3643 | 15 | 14.9879 | 0.8708 | 234.9412 | 18 | 17.9881 | 1.0157 | 269.2882 | 21 | 20.9877 | 1.1599 | 299.2785 | 24 | 23.9896 | 1.3106 | 326.7712 | 12 | 11.988 | 0.7008 | 178.1738 | 15 | 14.9879 | 0.8708 | 214.8053 | 18 | 17.9881 | 1.0157 | 247.2201 | 21 | 20.9877 | 1.1599 | 275.3368 | 24 | 23.9896 | 1.3106 | 302.1283 | 12 | 11.988 | 0.7008 | 183.1393 | 15 | 14.9879 | 0.8708 | 221.0287 | 18 | 17.9881 | 1.0157 | 253.5249 | 21 | 20.9877 | 1.1599 | 281.7327 | 24 | 23.9896 | 1.3106 | 307.8247 | 12 | 11.988 | 0.7008 | 190.3289 | 15 | 14.9879 | 0.8708 | 231.0309 | 18 | 17.9881 | 1.0157 | 266.112 | 21 | 20.9877 | 1.1599 | 294.1396 | 24 | 23.9896 | 1.3106 | 321.1835 | 12 | 11.988 | 0.7008 | 182.5719 | 15 | 14.9879 | 0.8708 | 220.5263 | 18 | 17.9881 | 1.0157 | 252.6007 | 21 | 20.9877 | 1.1599 | 281.163 | 24 | 23.9896 | 1.3106 | 307.1374 | 12 | 11.988 | 0.7008 | 188.7338 | 15 | 14.9879 | 0.8708 | 227.434 | 18 | 17.9881 | 1.0157 | 260.4196 | 21 | 20.9877 | 1.1599 | 289.8592 | 24 | 23.9896 | 1.3106 | 316.9514 | 18.1458 | 18.8391 | 17.8677 | 18.2898 | 1.7161 | 1.7869 | 3.7841 | 1.9654 | 11762.29379Mohm | 58.9601 | 31.9323 | 111.91215kohm | 267.3416 | 53.9161 | 16661.83407Mohm | 52.9189 | 31.9662 | 110.77213kohm | 19.1153 | 1 | 9092.273Mohm | 65.9659 | 33.0692 | 111.66388kohm | 35.3191 | 40.8439 | 9525.14329Mohm | 58.0177 | 32.6727 | 109.68406kohm | 26.8716 | 1 | 18176.81784Mohm | 67.0374 | 32.7947 | 110.70831kohm | 48.7399 | 37.1314 | 0.0Mohm | 62.1543 | 32.4667 | 111.19799kohm | 66.0825 | 50.0832 | 14285.64273Mohm | 65.1194 | 33.1334 | 112.1429kohm | 51.6525 | 30.7201 | 100018.0015Mohm | 66.6743 | 32.58 | 111.60083kohm | 142.9846 | 23.5371 | 66655.33596Mohm | 57.9843 | 33.1169 | 109.40435kohm | 7.1082 | 1 | 39993.40028Mohm | 64.2867 | 33.1969 | 109.28902kohm | 23.6744 | 1 | 25001.12513Mohm | 61.0777 | 32.2066 | 104.47212kohm | 0 | 1.9214 | 50007.50234Mohm | 59.2349 | 32.7194 | 110.01235kohm | 22.1019 | 24.2873 | 28568.715Mohm | 55.7764 | 32.9623 | 110.66232kohm | 25.5201 | 41.3127 | 12496.37555Mohm | 68.7519 | 32.4798 | 105.86734kohm | 24.862 | 1 | 28570.85646Mohm | 58.7386 | 33.3339 | 113.47936kohm | 263.8138 | 23.8386 | 22224.99912Mohm | 64.8436 | 32.8902 | 108.10811kohm | 38.2423 | 0.07007
002_2300V_test | 06/25/2026 | 05:21:39 PM | 23.9897 | 1.3131 | 326.7829 | 301.8367 | 307.6066 | 321.4504 | 307.4539 | 316.8798 | 12 | 11.9876 | 0.7026 | 193.338 | 15 | 14.9877 | 0.8702 | 234.866 | 18 | 17.988 | 1.0163 | 269.3033 | 21 | 20.9875 | 1.1599 | 299.3092 | 24 | 23.9895 | 1.311 | 326.7996 | 12 | 11.9876 | 0.7026 | 178.3407 | 15 | 14.9877 | 0.8702 | 215.6881 | 18 | 17.988 | 1.0163 | 247.7224 | 21 | 20.9875 | 1.1599 | 276.5568 | 24 | 23.9895 | 1.311 | 302.1905 | 12 | 11.9876 | 0.7026 | 183.2491 | 15 | 14.9877 | 0.8702 | 220.7791 | 18 | 17.988 | 1.0163 | 253.0327 | 21 | 20.9875 | 1.1599 | 281.6831 | 24 | 23.9895 | 1.311 | 308.0231 | 12 | 11.9876 | 0.7026 | 190.1275 | 15 | 14.9877 | 0.8702 | 231.0716 | 18 | 17.988 | 1.0163 | 264.5556 | 21 | 20.9875 | 1.1599 | 294.9311 | 24 | 23.9895 | 1.311 | 322.003 | 12 | 11.9876 | 0.7026 | 182.5438 | 15 | 14.9877 | 0.8702 | 220.511 | 18 | 17.988 | 1.0163 | 252.5926 | 21 | 20.9875 | 1.1599 | 281.0465 | 24 | 23.9895 | 1.311 | 307.4482 | 12 | 11.9876 | 0.7026 | 189.057 | 15 | 14.9877 | 0.8702 | 227.3337 | 18 | 17.988 | 1.0163 | 260.6475 | 21 | 20.9875 | 1.1599 | 290.1083 | 24 | 23.9895 | 1.311 | 317.0492 | 18.1496 | 18.8445 | 17.8719 | 18.2929 | 1.7186 | 1.788 | 3.7723 | 1.9674 | 12775.38876Mohm | 59.3501 | 31.6753 | 112.99091kohm | 21.1899 | 0 | 16423.7862Mohm | 55.3944 | 31.7158 | 110.17984kohm | 64.0046 | 1 | 9583.75062Mohm | 58.5927 | 32.893 | 110.36681kohm | 115.4204 | 29.062 | 10001.60902Mohm | 55.0698 | 32.443 | 108.26473kohm | 0 | 27.2883 | 19161.83311Mohm | 66.939 | 32.5242 | 111.20406kohm | 26.4122 | 33.8325 | 0.0Mohm | 66.6 | 32.1759 | 111.32202kohm | 33.7124 | 48.6043 | 14374.81264Mohm | 68.9776 | 32.9111 | 111.60885kohm | 80.6072 | 55.5237 | 115020.49818Mohm | 63.1081 | 32.3133 | 113.14087kohm | 0 | 0.04174 | 57490.50421Mohm | 36.0919 | 32.8419 | 109.9757kohm | 28.3944 | 0 | 32851.42983Mohm | 63.8015 | 32.8502 | 110.3418kohm | 63.2241 | 1 | 20910.001Mohm | 43.3105 | 31.908 | 103.53336kohm | 19.6097 | 0.10254 | 38339.33598Mohm | 57.2104 | 32.3125 | 110.12032kohm | 36.3164 | 0 | 25552.77676Mohm | 58.3047 | 32.5562 | 111.17868kohm | 26.43 | 62.5682 | 13525.82276Mohm | 62.9642 | 32.0483 | 105.82686kohm | 26.0489 | 1 | 22999.29897Mohm | 5.4269 | 32.9068 | 113.41751kohm | 62.7231 | 26.2388 | 23002.49966Mohm | 59.9124 | 32.4452 | 109.4054kohm | 70.3025 | 0.10223
0002_amshi_1500_coatedboards | 06/30/2026 | 09:48:58 AM | 23.9905 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 12 | 11.9895 | 0 | 0 | 15 | 14.9892 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9887 | 0 | 0 | 24 | 23.9906 | 0 | 0 | 18.1109 | 18.8088 | 17.8265 | 18.2523 | 1.7246 | 1.7962 | 3.7885 | 1.9671
0002_amshi_1500_coatedpart2 | 06/30/2026 | 10:16:43 AM | 23.9903 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 12 | 11.9891 | 0 | 0 | 15 | 14.9889 | 0 | 0 | 18 | 17.9893 | 0 | 0 | 21 | 20.9884 | 0 | 0 | 24 | 23.9902 | 0 | 0 | 18.1224 | 18.8198 | 17.8371 | 18.2638 | 1.7124 | 1.769 | 3.8748 | 1.9681 | 14.2443Mohm | 0 | 24.6436 | 112.01199kohm | 0 | 0.05607 | 31.01821Mohm | 86.7102 | 6.4891 | 107.95234kohm | 0 | 1 | 10002.5997Mohm | 68.7057 | 32.4856 | 110.24913kohm | 0 | 0 | 8823.8826Mohm | 43.7597 | 32.9626 | 108.61281kohm | 82.775 | 23.9578 | 12.95095Mohm | 34.4538 | 19.5329 | 109.75542kohm | 120.3498 | 33.383 | 93.52963Mohm | 19.5198 | 9.7995 | 110.01943kohm | 0 | 54.0461 | 42.3339Mohm | -47.1532 | 110.6222 | 107.08428kohm | 48.5625 | 18.9036 | 136.99269Mohm | 0.00852 | 11.0171 | 111.87626kohm | 74.8224 | 15.855 | 21425.99914Mohm | 94.4593 | 32.6089 | 109.29789kohm | 20.4586 | 0.03023 | 12495.58337Mohm | 49.8996 | 32.8258 | 108.58753kohm | 39.9253 | 1 | 16670.33218Mohm | 82.9577 | 31.7792 | 104.24628kohm | 0 | 1 | 15003.39943Mohm | 49.9494 | 32.3253 | 110.13287kohm | 49.5663 | 1 | 14996.70022Mohm | 95.0976 | 32.4371 | 111.45874kohm | 8.1252 | 82.4915 | 5553.55519Mohm | 57.1029 | 32.1981 | 105.67569kohm | 239.3471 | 1 | 12501.58312Mohm | 100.7903 | 32.6823 | 114.18028kohm | 3.1185 | 0.05176 | 15002.29958Mohm | 51.4374 | 32.4721 | 109.40551kohm | 932.829 | 0.2842
0002_amshi_1500_dry | 06/30/2026 | 02:44:49 PM | 23.9905 | 1.2958 | 316.4781 | 314.4966 | 314.0216 | 305.5956 | 311.9799 | 318.4865 | 12 | 11.9889 | 0.6993 | 188.025 | 15 | 14.9889 | 0.863 | 228.0368 | 18 | 17.9888 | 1.0065 | 260.7782 | 21 | 20.9883 | 1.1459 | 289.4904 | 24 | 23.9903 | 1.2939 | 316.2529 | 12 | 11.9889 | 0.6993 | 188.1752 | 15 | 14.9889 | 0.863 | 226.9952 | 18 | 17.9888 | 1.0065 | 259.416 | 21 | 20.9883 | 1.1459 | 287.9866 | 24 | 23.9903 | 1.2939 | 314.4951 | 12 | 11.9889 | 0.6993 | 185.324 | 15 | 14.9889 | 0.863 | 224.846 | 18 | 17.9888 | 1.0065 | 258.1672 | 21 | 20.9883 | 1.1459 | 287.4984 | 24 | 23.9903 | 1.2939 | 314.2574 | 12 | 11.9889 | 0.6993 | 182.4161 | 15 | 14.9889 | 0.863 | 220.0833 | 18 | 17.9888 | 1.0065 | 251.5663 | 21 | 20.9883 | 1.1459 | 279.989 | 24 | 23.9903 | 1.2939 | 305.0754 | 12 | 11.9889 | 0.6993 | 182.2507 | 15 | 14.9889 | 0.863 | 221.2088 | 18 | 17.9888 | 1.0065 | 254.5925 | 21 | 20.9883 | 1.1459 | 284.4819 | 24 | 23.9903 | 1.2939 | 312.1677 | 12 | 11.9889 | 0.6993 | 187.0817 | 15 | 14.9889 | 0.863 | 227.7296 | 18 | 17.9888 | 1.0065 | 262.0903 | 21 | 20.9883 | 1.1459 | 291.9998 | 24 | 23.9903 | 1.2939 | 319.2834 | 18.1341 | 18.8481 | 17.8664 | 18.2746 | 1.7114 | 1.7768 | 3.975 | 1.9692 | 11535.30785Mohm | 47.5135 | 31.9592 | 112.35202kohm | 40.7357 | 0.05101 | 11534.46181Mohm | 47.4822 | 33.2805 | 109.56842kohm | 36.9238 | 1 | 8825.47024Mohm | 69.7147 | 32.346 | 111.42701kohm | 0 | 0 | 8333.16617Mohm | 49.811 | 33.1298 | 110.29932kohm | 0 | 1 | 9372.93798Mohm | 52.1233 | 32.7944 | 110.99837kohm | 0 | 55.6848 | 12497.0004Mohm | 53.3241 | 33.209 | 110.05534kohm | 88.8665 | 33.8881 | 10000.8663Mohm | 49.3327 | 33.518 | 111.14777kohm | 59.3856 | 33.4071 | 13636.8187Mohm | 56.1483 | 33.1901 | 112.23486kohm | 46.2768 | 24.2353 | 16664.55554Mohm | 41.1433 | 33.0115 | 108.56759kohm | 35.9582 | 1 | 9371.31292Mohm | 51.3522 | 32.8712 | 109.56705kohm | 55.1321 | 1 | 11540.77003Mohm | 46.3631 | 32.1276 | 104.423kohm | 8.4457 | 1 | 10717.14299Mohm | 52.318 | 32.3516 | 110.81066kohm | 27.3435 | 35.4837 | 10711.85734Mohm | 40.4046 | 32.7294 | 107.66913kohm | 455.9422 | 44.1579 | 4837.19361Mohm | 60.5422 | 32.2189 | 106.77721kohm | 28.6752 | 1 | 10715.71389Mohm | 45.8093 | 33.06 | 114.22998kohm | 21.8123 | 29.3055 | 10001.9332Mohm | 57.9049 | 32.4973 | 108.22558kohm | 25.7196 | 24.0361
0002_amshi_1800_cb1 | 06/30/2026 | 03:15:08 PM | 23.9907 | 1.2941 | 317.1839 | 315.1879 | 314.9027 | 306.6399 | 312.2417 | 319.45 | 12 | 11.9889 | 0.6964 | 187.7308 | 15 | 14.989 | 0.8627 | 227.7722 | 18 | 17.989 | 1.0053 | 261.0911 | 21 | 20.9885 | 1.145 | 290.2502 | 24 | 23.9903 | 1.2919 | 317.0453 | 12 | 11.9889 | 0.6964 | 188.776 | 15 | 14.989 | 0.8627 | 227.4947 | 18 | 17.989 | 1.0053 | 259.943 | 21 | 20.9885 | 1.145 | 288.4751 | 24 | 23.9903 | 1.2919 | 315.1602 | 12 | 11.9889 | 0.6964 | 186.1316 | 15 | 14.989 | 0.8627 | 225.5547 | 18 | 17.989 | 1.0053 | 258.8968 | 21 | 20.9885 | 1.145 | 288.2731 | 24 | 23.9903 | 1.2919 | 315.2779 | 12 | 11.9889 | 0.6964 | 183.2533 | 15 | 14.989 | 0.8627 | 220.887 | 18 | 17.989 | 1.0053 | 252.6193 | 21 | 20.9885 | 1.145 | 280.4488 | 24 | 23.9903 | 1.2919 | 306.3978 | 12 | 11.9889 | 0.6964 | 182.7579 | 15 | 14.989 | 0.8627 | 221.8506 | 18 | 17.989 | 1.0053 | 255.1337 | 21 | 20.9885 | 1.145 | 284.9603 | 24 | 23.9903 | 1.2919 | 312.8317 | 12 | 11.9889 | 0.6964 | 187.2101 | 15 | 14.989 | 0.8627 | 227.6843 | 18 | 17.989 | 1.0053 | 262.0337 | 21 | 20.9885 | 1.145 | 291.9943 | 24 | 23.9903 | 1.2919 | 319.317 | 18.135 | 18.8509 | 17.869 | 18.2755 | 1.7193 | 1.7966 | 3.9556 | 1.9655 | 14996.49997Mohm | 45.947 | 32.2054 | 112.27679kohm | 17.3581 | 0.05228 | 12852.49977Mohm | 45.1995 | 32.45 | 110.05131kohm | 94.9903 | 1 | 11251.18739Mohm | 47.0868 | 32.5451 | 111.23228kohm | 265.4092 | 0.04104 | 9000.09996Mohm | 47.8121 | 32.3807 | 109.70624kohm | 21.9543 | 23.1977 | 11247.00037Mohm | 46.295 | 32.6646 | 110.49226kohm | 30.9239 | 60.8055 | 14994.41664Mohm | 52.4611 | 32.423 | 110.7071kohm | 25.3624 | 30.711 | 12856.78531Mohm | 47.7961 | 33.2085 | 111.98462kohm | 0 | 37.8734 | 18001.09959Mohm | 50.8146 | 32.4245 | 112.06473kohm | 0 | 0 | 29994.99986Mohm | 49.2072 | 32.7221 | 109.56376kohm | 0 | 1 | 11995.99969Mohm | 49.6866 | 32.3635 | 109.52302kohm | 33.654 | 1 | 16364.90931Mohm | 47.553 | 31.9433 | 104.46736kohm | 7.9834 | 1 | 13847.92343Mohm | 50.5739 | 31.9372 | 110.61828kohm | 0 | 0 | 16360.90874Mohm | 48.4323 | 32.4441 | 110.2507kohm | 59.0536 | 26.7309 | 6204.62075Mohm | 53.2477 | 31.7053 | 107.17464kohm | 39.8771 | 1 | 15000.58322Mohm | 48.5737 | 32.7567 | 113.71744kohm | 55.9905 | 24.2158 | 13847.38539Mohm | 53.0616 | 31.9778 | 107.66809kohm | 22.9069 | 1
0002_amshi_2000_cb1 | 06/30/2026 | 04:40:23 PM | 23.9874 | 1.2999 | 317.7097 | 315.4534 | 315.4556 | 306.249 | 313.0979 | 319.6992 | 12 | 11.9846 | 0.6911 | 188.421 | 15 | 14.9851 | 0.8603 | 227.762 | 18 | 17.9857 | 1.0069 | 261.2778 | 21 | 20.9857 | 1.1506 | 290.3025 | 24 | 23.988 | 1.3029 | 317.1046 | 12 | 11.9846 | 0.6911 | 187.8753 | 15 | 14.9851 | 0.8603 | 227.1395 | 18 | 17.9857 | 1.0069 | 259.6912 | 21 | 20.9857 | 1.1506 | 288.3906 | 24 | 23.988 | 1.3029 | 315.42 | 12 | 11.9846 | 0.6911 | 185.7762 | 15 | 14.9851 | 0.8603 | 225.6098 | 18 | 17.9857 | 1.0069 | 258.5217 | 21 | 20.9857 | 1.1506 | 288.1424 | 24 | 23.988 | 1.3029 | 316.0058 | 12 | 11.9846 | 0.6911 | 182.5571 | 15 | 14.9851 | 0.8603 | 220.5452 | 18 | 17.9857 | 1.0069 | 251.8899 | 21 | 20.9857 | 1.1506 | 280.295 | 24 | 23.988 | 1.3029 | 306.1667 | 12 | 11.9846 | 0.6911 | 182.9043 | 15 | 14.9851 | 0.8603 | 220.8901 | 18 | 17.9857 | 1.0069 | 253.4637 | 21 | 20.9857 | 1.1506 | 283.0847 | 24 | 23.988 | 1.3029 | 309.9584 | 12 | 11.9846 | 0.6911 | 187.3784 | 15 | 14.9851 | 0.8603 | 227.6136 | 18 | 17.9857 | 1.0069 | 261.6608 | 21 | 20.9857 | 1.1506 | 291.4794 | 24 | 23.988 | 1.3029 | 318.6602 | 18.1536 | 18.8672 | 17.8819 | 18.2855 | 1.706 | 1.7814 | 3.961 | 1.9447 | 15381.84683Mohm | 50.6652 | 33.1009 | 112.76893kohm | 1562.8829 | 0.05405 | 16661.50041Mohm | 55.3759 | 33.1516 | 108.92081kohm | 62.2923 | 1 | 11766.05858Mohm | 51.4734 | 33.4428 | 110.65664kohm | 33.6176 | 24.6062 | 10527.05284Mohm | 58.3629 | 33.0821 | 109.38011kohm | 216.1202 | 1 | 11761.35313Mohm | 54.8923 | 33.6413 | 110.66081kohm | 169.3634 | 25.7767 | 24993.37519Mohm | 70.3953 | 33.2411 | 111.0365kohm | 41.2277 | 29.2109 | 13332.6663Mohm | 57.142 | 33.9829 | 111.105kohm | 22.673 | 1127.5943 | 25003.87629Mohm | 64.2285 | 33.1983 | 111.80327kohm | 35.4043 | 24.3112 | 28567.14204Mohm | 55.9574 | 33.517 | 110.24106kohm | 29.0324 | 1 | 15381.84683Mohm | 65.0337 | 33.2383 | 109.27994kohm | 17.9671 | 1 | 16667.0007Mohm | 55.3018 | 32.6023 | 104.73723kohm | 0.0637 | 1 | 16667.667Mohm | 64.9952 | 32.7004 | 110.51636kohm | 98.3897 | 24.162 | 15382.53888Mohm | 50.156 | 33.2335 | 111.65671kohm | 6.5001 | 30.1161 | 7405.62941Mohm | 79.0556 | 32.4536 | 106.94204kohm | 200.4817 | 1 | 15383.84691Mohm | 54.0896 | 33.5786 | 113.83434kohm | 0 | 22.0449 | 16667.50017Mohm | 72.2883 | 32.8574 | 108.12277kohm | 28.6217 | 1
0002_amshi_2300_cb1 | 06/30/2026 | 05:32:20 PM | 23.9904 | 1.3069 | 315.2534 | 317.5092 | 314.0497 | 307.2904 | 308.8213 | 319.6279 | 12 | 11.9887 | 0.6943 | 186.4339 | 15 | 14.9887 | 0.8642 | 225.7788 | 18 | 17.9888 | 1.0069 | 259.0367 | 21 | 20.9883 | 1.1515 | 287.9598 | 24 | 23.9901 | 1.3026 | 315.3294 | 12 | 11.9887 | 0.6943 | 190.7535 | 15 | 14.9887 | 0.8642 | 229.5141 | 18 | 17.9888 | 1.0069 | 262.166 | 21 | 20.9883 | 1.1515 | 290.677 | 24 | 23.9901 | 1.3026 | 317.6476 | 12 | 11.9887 | 0.6943 | 185.8348 | 15 | 14.9887 | 0.8642 | 224.8473 | 18 | 17.9888 | 1.0069 | 257.7185 | 21 | 20.9883 | 1.1515 | 286.5802 | 24 | 23.9901 | 1.3026 | 313.3532 | 12 | 11.9887 | 0.6943 | 183.0405 | 15 | 14.9887 | 0.8642 | 220.8013 | 18 | 17.9888 | 1.0069 | 252.7364 | 21 | 20.9883 | 1.1515 | 281.1536 | 24 | 23.9901 | 1.3026 | 307.2641 | 12 | 11.9887 | 0.6943 | 182.4081 | 15 | 14.9887 | 0.8642 | 220.0227 | 18 | 17.9888 | 1.0069 | 252.3529 | 21 | 20.9883 | 1.1515 | 281.8973 | 24 | 23.9901 | 1.3026 | 308.6185 | 12 | 11.9887 | 0.6943 | 187.6594 | 15 | 14.9887 | 0.8642 | 227.7943 | 18 | 17.9888 | 1.0069 | 262.5473 | 21 | 20.9883 | 1.1515 | 292.4218 | 24 | 23.9901 | 1.3026 | 320.0068 | 18.142 | 18.844 | 17.878 | 18.2836 | 1.7016 | 1.7834 | 3.8341 | 1.9601 | 14372.81221Mohm | 54.0281 | 32.7998 | 111.22899kohm | 23.2767 | 123.973 | 19161.7497Mohm | 59.6395 | 32.85 | 109.94588kohm | 44.0449 | 1 | 11500.60041Mohm | 58.2264 | 33.1445 | 111.4747kohm | 83.0843 | 23.2545 | 11500.84944Mohm | 57.5266 | 32.7758 | 109.54873kohm | 34.228 | 23.7742 | 11497.49983Mohm | 61.4678 | 33.3376 | 110.83605kohm | 37.7979 | 81.9627 | 32847.8584Mohm | 68.7 | 32.9344 | 111.08626kohm | 75.5107 | 45.2623 | 12777.22117Mohm | 59.6607 | 33.6877 | 110.13457kohm | 30.0357 | 40.8385 | 8519.48118Mohm | 34.7542 | 32.8558 | 112.07603kohm | 31.7831 | 24.2763 | 25551.44213Mohm | 74.0646 | 33.1847 | 110.20407kohm | 26.0456 | 1 | 20905.18255Mohm | 63.5545 | 32.959 | 109.4442kohm | 119.1082 | 1 | 16428.35686Mohm | 43.9983 | 32.2544 | 103.96406kohm | 34.1704 | 1 | 20910.72676Mohm | 64.8867 | 32.4536 | 110.22813kohm | 26.2153 | 0.05564 | 16426.78564Mohm | 58.6679 | 32.94 | 111.5138kohm | 30.4515 | 49.0025 | 9197.91992Mohm | 75.0434 | 32.1937 | 107.57837kohm | 232.8936 | 1 | 16427.99937Mohm | 88.6497 | 33.2508 | 113.61549kohm | 87.0462 | 31.6163 | 20910.81776Mohm | 69.6741 | 32.6174 | 108.1873kohm | 34.0245 | 0.09722
005_amshi_1500_c03 | 07/01/2026 | 04:25:01 PM
005_amshi_1500_c03 | 07/01/2026 | 04:25:40 PM
005_amshi_1500_c03 | 07/01/2026 | 04:25:49 PM
005_amshi_1500_c03 | 07/01/2026 | 04:28:23 PM | 23.9914 | 1.3076 | 314.5055 | 317.118 | 312.5214 | 306.9689 | 308.1958 | 319.8897 | 12 | 11.9906 | 0.7019 | 185.3961 | 15 | 14.9903 | 0.8674 | 225.3082 | 18 | 17.9901 | 1.011 | 258.4964 | 21 | 20.9894 | 1.1525 | 287.2898 | 24 | 23.9911 | 1.302 | 314.7515 | 12 | 11.9906 | 0.7019 | 189.7663 | 15 | 14.9903 | 0.8674 | 229.309 | 18 | 17.9901 | 1.011 | 262.1639 | 21 | 20.9894 | 1.1525 | 290.7508 | 24 | 23.9911 | 1.302 | 317.1902 | 12 | 11.9906 | 0.7019 | 186.2214 | 15 | 14.9903 | 0.8674 | 223.6437 | 18 | 17.9901 | 1.011 | 257.3715 | 21 | 20.9894 | 1.1525 | 286.3704 | 24 | 23.9911 | 1.302 | 313.6494 | 12 | 11.9906 | 0.7019 | 182.5463 | 15 | 14.9903 | 0.8674 | 220.1464 | 18 | 17.9901 | 1.011 | 252.3025 | 21 | 20.9894 | 1.1525 | 280.7052 | 24 | 23.9911 | 1.302 | 307.2578 | 12 | 11.9906 | 0.7019 | 181.7836 | 15 | 14.9903 | 0.8674 | 219.7575 | 18 | 17.9901 | 1.011 | 252.1805 | 21 | 20.9894 | 1.1525 | 281.1391 | 24 | 23.9911 | 1.302 | 308.5961 | 12 | 11.9906 | 0.7019 | 187.3144 | 15 | 14.9903 | 0.8674 | 227.5954 | 18 | 17.9901 | 1.011 | 261.7024 | 21 | 20.9894 | 1.1525 | 291.8749 | 24 | 23.9911 | 1.302 | 319.4509 | 18.719 | 18.2152 | 17.8357 | 18.2542 | 1.7568 | 1.7364 | 3.9086 | 1.9541 | 16662.55495Mohm | 57.0945 | 33.6042 | 112.16571kohm | 1605.1623 | 78.3577 | 14994.79958Mohm | 44.0223 | 33.4134 | 107.7704kohm | 82.1076 | 1 | 10716.50037Mohm | 48.4188 | 33.853 | 111.26868kohm | 27.7031 | 0 | 9375.18789Mohm | 42.9131 | 33.2418 | 110.30495kohm | 0 | 23.7675 | 12495.83361Mohm | 51.8541 | 34.1345 | 110.53016kohm | 68.3408 | 32.0972 | 16659.77717Mohm | 49.3113 | 33.4547 | 111.23539kohm | 56.7385 | 43.6093 | 7895.36849Mohm | 112.4551 | 34.6211 | 110.55579kohm | 425.1592 | 69.2491 | 1339.2589Mohm | 0 | 33.3785 | 112.75688kohm | 96.0971 | 21.2843 | 24995.83389Mohm | 101.3333 | 32.5979 | 109.81204kohm | 75.6658 | 185107.1522 | 10710.42825Mohm | 45.6381 | 33.9675 | 109.14635kohm | 29.9997 | 1 | 15002.90017Mohm | 68.2797 | 32.0857 | 104.91629kohm | 0 | 1 | 12502.91674Mohm | 48.9679 | 33.8065 | 110.04888kohm | 5.6934 | 23.3809 | 13633.45511Mohm | 103.8418 | 33.3927 | 110.67614kohm | 42.7497 | 36.3085 | 4998.33313Mohm | 54.5346 | 34.2437 | 105.70648kohm | 64.4476 | 1 | 11539.84699Mohm | 111.8713 | 33.1605 | 114.43327kohm | 223.1424 | 16.9472 | 11539.84699Mohm | 53.9288 | 33.9842 | 109.16222kohm | 44.5774 | 24.626
0005_amshi_50V_safetest_firstwiecQC | 07/10/2026 | 02:15:37 PM
05_amshi_wiecQC | 07/10/2026 | 02:21:56 PM | 23.9911 | 1.3008 | 313.451 | 316.7103 | 313.7709 | 306.9731 | 307.1168 | 317.612 | 12 | 11.9902 | 0.699 | 185.9575 | 15 | 14.9897 | 0.8641 | 225.6765 | 18 | 17.9896 | 1.0074 | 258.3421 | 21 | 20.989 | 1.1474 | 287.9942 | 24 | 23.9908 | 1.2947 | 314.3387 | 12 | 11.9902 | 0.699 | 189.884 | 15 | 14.9897 | 0.8641 | 228.8738 | 18 | 17.9896 | 1.0074 | 261.2511 | 21 | 20.989 | 1.1474 | 289.8083 | 24 | 23.9908 | 1.2947 | 316.4551 | 12 | 11.9902 | 0.699 | 185.2724 | 15 | 14.9897 | 0.8641 | 224.1957 | 18 | 17.9896 | 1.0074 | 257.3295 | 21 | 20.989 | 1.1474 | 286.837 | 24 | 23.9908 | 1.2947 | 313.1779 | 12 | 11.9902 | 0.699 | 182.1257 | 15 | 14.9897 | 0.8641 | 220.7223 | 18 | 17.9896 | 1.0074 | 252.4081 | 21 | 20.989 | 1.1474 | 280.8281 | 24 | 23.9908 | 1.2947 | 307.1996 | 12 | 11.9902 | 0.699 | 180.6737 | 15 | 14.9897 | 0.8641 | 218.7907 | 18 | 17.9896 | 1.0074 | 250.9603 | 21 | 20.989 | 1.1474 | 280.2377 | 24 | 23.9908 | 1.2947 | 307.2175 | 12 | 11.9902 | 0.699 | 186.1801 | 15 | 14.9897 | 0.8641 | 228.5778 | 18 | 17.9896 | 1.0074 | 262.6242 | 21 | 20.989 | 1.1474 | 292.4754 | 24 | 23.9908 | 1.2947 | 319.5176 | 18.7188 | 18.2142 | 17.8333 | 18.2522 | 1.7567 | 1.7408 | 3.8988 | 1.9641 | 261.63158Mohm | 12.0958 | 26.0573 | 111.53479kohm | 0 | 108.2692 | 311.75001Mohm | 19.6133 | 25.8479 | 108.94572kohm | 0 | 1 | 277.16665Mohm | 18.6938 | 26.1039 | 112.01976kohm | 35.3482 | 0.03781 | 248.7Mohm | 19.9989 | 25.755 | 107.18961kohm | 30.6545 | 24.3632 | 234.85715Mohm | 20.8403 | 26.1546 | 110.67172kohm | 34.9591 | 62.4751 | 329.66666Mohm | 17.7329 | 25.7252 | 110.51515kohm | 24.3291 | 63.05 | 238.19049Mohm | 10.4397 | 26.2766 | 110.31772kohm | 50.5833 | 37.5067 | 230.09091Mohm | 12.9146 | 25.8389 | 111.05045kohm | 0.05154 | 0.04508 | 447.54545Mohm | 18.4948 | 26.0812 | 110.4481kohm | 59.8788 | 1 | 225.5Mohm | 19.1476 | 25.8302 | 108.36481kohm | 0 | 1 | 248.09999Mohm | 0.04888 | 26.0332 | 104.41489kohm | 0 | 10.6255 | 238.33334Mohm | 12.3012 | 25.9069 | 110.4398kohm | 18.144 | 23.2793 | 247.64999Mohm | 17.6134 | 26.5048 | 111.52939kohm | 78.339 | 56.9218 | 118.92683Mohm | 19.798 | 26.0644 | 105.62353kohm | 13.4308 | 1 | 230.13637Mohm | 0.0346 | 26.3079 | 113.67605kohm | 0 | 0 | 236.4762Mohm | 15.6046 | 26.0048 | 109.49825kohm | 102.6015 | 1
05_amshi_wiecqc | 07/10/2026 | 02:44:33 PM | 23.9907 | 1.295 | 314.6991 | 317.7987 | 314.2648 | 307.7224 | 307.7428 | 319.9994 | 12 | 11.9895 | 0.6924 | 186.0353 | 15 | 14.9894 | 0.8618 | 226.5334 | 18 | 17.9894 | 1.0045 | 259.5004 | 21 | 20.9887 | 1.1457 | 288.0045 | 24 | 23.9904 | 1.2909 | 315.0749 | 12 | 11.9895 | 0.6924 | 190.5432 | 15 | 14.9894 | 0.8618 | 229.5278 | 18 | 17.9894 | 1.0045 | 262.4201 | 21 | 20.9887 | 1.1457 | 291.2787 | 24 | 23.9904 | 1.2909 | 317.8881 | 12 | 11.9895 | 0.6924 | 186.1897 | 15 | 14.9894 | 0.8618 | 225.1419 | 18 | 17.9894 | 1.0045 | 258.4871 | 21 | 20.9887 | 1.1457 | 287.5248 | 24 | 23.9904 | 1.2909 | 314.0208 | 12 | 11.9895 | 0.6924 | 183.4138 | 15 | 14.9894 | 0.8618 | 221.546 | 18 | 17.9894 | 1.0045 | 253.6164 | 21 | 20.9887 | 1.1457 | 281.4213 | 24 | 23.9904 | 1.2909 | 307.4209 | 12 | 11.9895 | 0.6924 | 181.8577 | 15 | 14.9894 | 0.8618 | 219.6503 | 18 | 17.9894 | 1.0045 | 252.0352 | 21 | 20.9887 | 1.1457 | 281.1868 | 24 | 23.9904 | 1.2909 | 308.2498 | 12 | 11.9895 | 0.6924 | 188.2076 | 15 | 14.9894 | 0.8618 | 229.2293 | 18 | 17.9894 | 1.0045 | 263.6802 | 21 | 20.9887 | 1.1457 | 293.432 | 24 | 23.9904 | 1.2909 | 320.4497 | 18.7329 | 18.2238 | 17.8472 | 18.2624 | 1.7651 | 1.7525 | 3.8974 | 1.9725
5_amshi_wiecquick | 07/10/2026 | 02:55:13 PM | 23.9905 | 1.2978 | 314.7461 | 317.3254 | 314.1193 | 307.9211 | 307.8439 | 320.4245 | 12 | 11.9889 | 0.6949 | 186.4896 | 15 | 14.989 | 0.8629 | 225.4029 | 18 | 17.9891 | 1.0055 | 258.9085 | 21 | 20.9884 | 1.1452 | 287.9659 | 24 | 23.9902 | 1.2949 | 314.8662 | 12 | 11.9889 | 0.6949 | 189.778 | 15 | 14.989 | 0.8629 | 228.4788 | 18 | 17.9891 | 1.0055 | 261.4252 | 21 | 20.9884 | 1.1452 | 290.403 | 24 | 23.9902 | 1.2949 | 316.4155 | 12 | 11.9889 | 0.6949 | 186.1882 | 15 | 14.989 | 0.8629 | 225.0528 | 18 | 17.9891 | 1.0055 | 258.275 | 21 | 20.9884 | 1.1452 | 287.1539 | 24 | 23.9902 | 1.2949 | 314.1034 | 12 | 11.9889 | 0.6949 | 182.9119 | 15 | 14.989 | 0.8629 | 221.1198 | 18 | 17.9891 | 1.0055 | 253.0524 | 21 | 20.9884 | 1.1452 | 281.6262 | 24 | 23.9902 | 1.2949 | 307.5308 | 12 | 11.9889 | 0.6949 | 181.3678 | 15 | 14.989 | 0.8629 | 219.2277 | 18 | 17.9891 | 1.0055 | 251.8334 | 21 | 20.9884 | 1.1452 | 280.8149 | 24 | 23.9902 | 1.2949 | 307.858 | 12 | 11.9889 | 0.6949 | 187.9486 | 15 | 14.989 | 0.8629 | 228.9781 | 18 | 17.9891 | 1.0055 | 262.8036 | 21 | 20.9884 | 1.1452 | 293.0358 | 24 | 23.9902 | 1.2949 | 320.0783 | 18.7326 | 18.227 | 17.8471 | 18.2653 | 1.7616 | 1.7331 | 3.9605 | 1.9702
05_amshi_wiec_test | 07/10/2026 | 03:06:53 PM | 23.9905 | 1.3004 | 315.1443 | 316.7225 | 314.3976 | 308.026 | 307.1867 | 320.1894
5_amshi_wiec | 07/10/2026 | 03:21:50 PM | 23.9906 | 1.2982 | 314.7781 | 317.4658 | 314.5258 | 307.4909 | 307.8105 | 320.2185
5_amshi_wiec | 07/10/2026 | 03:46:34 PM | 23.9899 | 1.3007 | 314.9929 | 317.2318 | 314.1324 | 307.7925 | 307.4499 | 319.1113
0005_amshi_1500_V_test | 07/16/2026 | 10:13:42 AM | 23.9908 | 1.301 | 314.8579 | 318.1321 | 313.4295 | 307.7683 | 310.419 | 319.1862 | 12 | 11.9906 | 0.692 | 186.4034 | 15 | 14.99 | 0.8613 | 226.2471 | 18 | 17.9895 | 1.0028 | 259.5256 | 21 | 20.9887 | 1.1445 | 288.6971 | 24 | 23.9903 | 1.2946 | 315.2419 | 12 | 11.9906 | 0.692 | 190.8766 | 15 | 14.99 | 0.8613 | 229.855 | 18 | 17.9895 | 1.0028 | 262.6314 | 21 | 20.9887 | 1.1445 | 292.2144 | 24 | 23.9903 | 1.2946 | 317.9759 | 12 | 11.9906 | 0.692 | 185.5727 | 15 | 14.99 | 0.8613 | 224.5098 | 18 | 17.9895 | 1.0028 | 257.6828 | 21 | 20.9887 | 1.1445 | 286.788 | 24 | 23.9903 | 1.2946 | 313.5194 | 12 | 11.9906 | 0.692 | 182.9985 | 15 | 14.99 | 0.8613 | 220.8645 | 18 | 17.9895 | 1.0028 | 253.0441 | 21 | 20.9887 | 1.1445 | 281.8962 | 24 | 23.9903 | 1.2946 | 307.8982 | 12 | 11.9906 | 0.692 | 182.9386 | 15 | 14.99 | 0.8613 | 221.1858 | 18 | 17.9895 | 1.0028 | 253.7552 | 21 | 20.9887 | 1.1445 | 283.5075 | 24 | 23.9903 | 1.2946 | 310.7378 | 12 | 11.9906 | 0.692 | 187.7536 | 15 | 14.99 | 0.8613 | 229.1018 | 18 | 17.9895 | 1.0028 | 263.4899 | 21 | 20.9887 | 1.1445 | 293.4057 | 24 | 23.9903 | 1.2946 | 320.7833 | 18.732 | 18.2678 | 17.858 | 18.2724 | 1.7639 | 1.7414 | 4.1534 | 1.9717 | 12496.83357Mohm | 0 | 35.7643 | 112.00135kohm | 0.0026 | 0 | 12495.58337Mohm | 0 | 34.2042 | 110.01339kohm | 0.00519 | 0.86602 | 11540.23104Mohm | 0 | 35.9971 | 111.26313kohm | 0.01032 | 0 | 9373.50027Mohm | 0 | 34.1048 | 110.40768kohm | 0.00208 | 0.86602 | 11535.61585Mohm | 0 | 36.2458 | 111.02967kohm | 0.00208 | 0.95592 | 14993.79983Mohm | 0 | 34.9689 | 110.5328kohm | 0.00519 | 0.95592 | 10001.13323Mohm | 0 | 36.3559 | 110.73122kohm | 0.0013 | 0.95592 | 10000.06633Mohm | 0 | 34.1549 | 112.48293kohm | 0.00346 | 0 | 24995.3334Mohm | 0 | 35.8081 | 108.90589kohm | 0 | 0.86602 | 7891.0003Mohm | 0 | 34.4966 | 109.38106kohm | 0.0013 | 0.86602 | 12503.16699Mohm | 0 | 35.2633 | 103.62997kohm | 0.0026 | 0.86602 | 10001.39934Mohm | 0 | 34.3266 | 110.2358kohm | 0.00116 | 0.86602 | 11535.61585Mohm | 0 | 36.5433 | 110.84933kohm | 0.01032 | 0 | 4283.97154Mohm | 0 | 34.8802 | 107.57083kohm | 0.01032 | 0.86602 | 10000.99977Mohm | 0 | 36.0729 | 113.91024kohm | 0.00174 | 0 | 8824.7055Mohm | 0 | 34.361 | 108.08594kohm | 0.0026 | 0.86602
0005_amshi_1800_V_test | 07/16/2026 | 10:59:49 AM | 23.9901 | 1.2982 | 314.8929 | 318.0536 | 313.6291 | 307.5174 | 309.9552 | 320.6406 | 12 | 11.9886 | 0.6885 | 186.9944 | 15 | 14.9886 | 0.8591 | 226.1536 | 18 | 17.9887 | 1.0033 | 259.2312 | 21 | 20.9879 | 1.1448 | 287.6997 | 24 | 23.9897 | 1.2945 | 314.6326 | 12 | 11.9886 | 0.6885 | 190.8229 | 15 | 14.9886 | 0.8591 | 229.7795 | 18 | 17.9887 | 1.0033 | 262.2927 | 21 | 20.9879 | 1.1448 | 291.1462 | 24 | 23.9897 | 1.2945 | 317.7703 | 12 | 11.9886 | 0.6885 | 185.7445 | 15 | 14.9886 | 0.8591 | 224.3485 | 18 | 17.9887 | 1.0033 | 257.5809 | 21 | 20.9879 | 1.1448 | 286.5515 | 24 | 23.9897 | 1.2945 | 313.2113 | 12 | 11.9886 | 0.6885 | 182.5582 | 15 | 14.9886 | 0.8591 | 220.8324 | 18 | 17.9887 | 1.0033 | 252.825 | 21 | 20.9879 | 1.1448 | 281.4025 | 24 | 23.9897 | 1.2945 | 307.5525 | 12 | 11.9886 | 0.6885 | 183.1983 | 15 | 14.9886 | 0.8591 | 221.35 | 18 | 17.9887 | 1.0033 | 254.0787 | 21 | 20.9879 | 1.1448 | 283.4452 | 24 | 23.9897 | 1.2945 | 310.8659 | 12 | 11.9886 | 0.6885 | 188.2197 | 15 | 14.9886 | 0.8591 | 228.7755 | 18 | 17.9887 | 1.0033 | 262.8341 | 21 | 20.9879 | 1.1448 | 292.9921 | 24 | 23.9897 | 1.2945 | 320.2422 | 18.7431 | 18.2826 | 17.8706 | 18.2869 | 1.7588 | 1.7346 | 4.08 | 1.9712 | 12854.49998Mohm | 0 | 35.979 | 111.04271kohm | 0.00112 | 0.94391 | 14994.08399Mohm | 0 | 34.6204 | 109.8037kohm | 0.00056 | 0.84463 | 12001.53273Mohm | 0 | 36.341 | 111.269kohm | 0.00074 | 0 | 10587.76485Mohm | 0 | 34.5397 | 108.14194kohm | 0.00149 | 0.84463 | 11246.75013Mohm | 0 | 35.8162 | 111.16112kohm | 0.00064 | 1e-05 | 29989.33376Mohm | 0 | 34.6923 | 109.19765kohm | 0.00223 | 1e-05 | 11249.50053Mohm | 0 | 36.5551 | 110.73142kohm | 0.00089 | 0.94902 | 15001.25054Mohm | 0 | 33.796 | 112.07556kohm | 0.00074 | 0 | 29994.16774Mohm | 0 | 36.1112 | 108.72611kohm | 0.95592 | 0.84463 | 11995.73276Mohm | 0 | 34.8884 | 108.82434kohm | 0.00074 | 0.84463 | 13846.6924Mohm | 0 | 35.3388 | 103.23983kohm | 0.00112 | 0.84463 | 12858.49953Mohm | 0 | 33.8947 | 109.95219kohm | 0.00223 | 0 | 12854.64298Mohm | 0 | 36.8298 | 109.82173kohm | 0.00064 | 1e-05 | 5623.06264Mohm | 0 | 34.5527 | 107.32281kohm | 0.00064 | 0.84463 | 12000.5993Mohm | 0 | 35.6422 | 114.45652kohm | 0.00112 | 0 | 12000.73276Mohm | 0 | 34.6568 | 109.37585kohm | 0.00056 | 0.84463
0005_amshi_2000_V_test | 07/16/2026 | 12:27:55 PM | 23.9901 | 1.2971 | 315.098 | 318.1211 | 313.357 | 307.8601 | 310.4869 | 320.0418 | 12 | 11.9884 | 0.6876 | 186.4017 | 15 | 14.9884 | 0.8585 | 226.1969 | 18 | 17.9885 | 1.0017 | 259.0279 | 21 | 20.9878 | 1.1448 | 287.9235 | 24 | 23.9899 | 1.2943 | 315.0704 | 12 | 11.9884 | 0.6876 | 190.7632 | 15 | 14.9884 | 0.8585 | 230.0679 | 18 | 17.9885 | 1.0017 | 262.3481 | 21 | 20.9878 | 1.1448 | 291.2543 | 24 | 23.9899 | 1.2943 | 318.1517 | 12 | 11.9884 | 0.6876 | 185.372 | 15 | 14.9884 | 0.8585 | 224.5895 | 18 | 17.9885 | 1.0017 | 257.6052 | 21 | 20.9878 | 1.1448 | 286.6216 | 24 | 23.9899 | 1.2943 | 313.3095 | 12 | 11.9884 | 0.6876 | 182.8513 | 15 | 14.9884 | 0.8585 | 221.0733 | 18 | 17.9885 | 1.0017 | 252.8905 | 21 | 20.9878 | 1.1448 | 281.293 | 24 | 23.9899 | 1.2943 | 307.9446 | 12 | 11.9884 | 0.6876 | 183.1864 | 15 | 14.9884 | 0.8585 | 221.263 | 18 | 17.9885 | 1.0017 | 254.1154 | 21 | 20.9878 | 1.1448 | 283.3274 | 24 | 23.9899 | 1.2943 | 310.3667 | 12 | 11.9884 | 0.6876 | 188.0543 | 15 | 14.9884 | 0.8585 | 229.0019 | 18 | 17.9885 | 1.0017 | 263.1032 | 21 | 20.9878 | 1.1448 | 292.7671 | 24 | 23.9899 | 1.2943 | 320.3409 | 18.7449 | 18.2836 | 17.8722 | 18.2885 | 1.759 | 1.7304 | 4.0138 | 1.9657 | 15381.61584Mohm | 0 | 35.6353 | 112.31621kohm | 0.0013 | 1e-05 | 16661.75066Mohm | 0 | 35.9403 | 109.78401kohm | 0.01031 | 0.84463 | 11765.9997Mohm | 0 | 36.3147 | 109.57034kohm | 0.00116 | 0 | 10526.89479Mohm | 0 | 35.8537 | 109.42266kohm | 0.0026 | 0 | 12496.62504Mohm | 0 | 36.6335 | 111.22977kohm | 0.00346 | 1e-05 | 49986.2499Mohm | 0 | 36.1168 | 110.86859kohm | 0.94902 | 0.94902 | 12499.06263Mohm | 0 | 36.6501 | 111.35599kohm | 0.00519 | 1e-05 | 16668.83379Mohm | 0 | 35.9372 | 111.77209kohm | 0.0026 | 1e-05 | 33328.16644Mohm | 0 | 36.0725 | 108.48556kohm | 0.00104 | 0.84463 | 15381.07779Mohm | 0 | 36.2242 | 109.14927kohm | 0.00174 | 0.84463 | 15384.46195Mohm | 0 | 35.1802 | 104.72562kohm | 0.00116 | 0.84463 | 15385.46199Mohm | 0 | 36.0614 | 111.35869kohm | 0.00149 | 0 | 14284.07151Mohm | 0 | 37.008 | 111.50255kohm | 0.00208 | 1e-05 | 6664.933Mohm | 0 | 36.6466 | 107.37691kohm | 0 | 0.84463 | 12499.43799Mohm | 0 | 35.3735 | 114.12601kohm | 0 | 0 | 16667.83383Mohm | 0 | 36.038 | 109.29738kohm | 0.00116 | 0
0005_amshi_2300_V_test | 07/16/2026 | 02:15:47 PM | 23.99 | 1.2963 | 315.1584 | 318.3872 | 313.7189 | 308.0372 | 310.7509 | 320.2092 | 12 | 11.9883 | 0.6917 | 186.2358 | 15 | 14.9883 | 0.8583 | 226.1476 | 18 | 17.9885 | 1.0008 | 259.1582 | 21 | 20.9879 | 1.1437 | 288.0319 | 24 | 23.9897 | 1.2934 | 314.528 | 12 | 11.9883 | 0.6917 | 190.7322 | 15 | 14.9883 | 0.8583 | 229.554 | 18 | 17.9885 | 1.0008 | 262.5386 | 21 | 20.9879 | 1.1437 | 291.2373 | 24 | 23.9897 | 1.2934 | 318.2578 | 12 | 11.9883 | 0.6917 | 185.7768 | 15 | 14.9883 | 0.8583 | 224.4908 | 18 | 17.9885 | 1.0008 | 258.0349 | 21 | 20.9879 | 1.1437 | 286.8878 | 24 | 23.9897 | 1.2934 | 313.4924 | 12 | 11.9883 | 0.6917 | 183.0224 | 15 | 14.9883 | 0.8583 | 221.5124 | 18 | 17.9885 | 1.0008 | 253.3915 | 21 | 20.9879 | 1.1437 | 282.1161 | 24 | 23.9897 | 1.2934 | 308.0078 | 12 | 11.9883 | 0.6917 | 183.0526 | 15 | 14.9883 | 0.8583 | 221.3246 | 18 | 17.9885 | 1.0008 | 254.3396 | 21 | 20.9879 | 1.1437 | 283.4586 | 24 | 23.9897 | 1.2934 | 311.0479 | 12 | 11.9883 | 0.6917 | 187.3385 | 15 | 14.9883 | 0.8583 | 228.8847 | 18 | 17.9885 | 1.0008 | 263.4332 | 21 | 20.9879 | 1.1437 | 293.1773 | 24 | 23.9897 | 1.2934 | 320.7407 | 18.7477 | 18.2861 | 17.874 | 18.2913 | 1.7626 | 1.7309 | 4.0668 | 1.9825 | 13527.52887Mohm | 0 | 36.1527 | 111.53932kohm | 0.0026 | 0.94902 | 19161.41807Mohm | 0 | 34.7145 | 109.80305kohm | 0.00174 | 0.84463 | 11499.99983Mohm | 0 | 36.6148 | 110.66827kohm | 0.00174 | 0 | 12106.05227Mohm | 0 | 34.6286 | 110.07128kohm | 0.00149 | 0.84463 | 10950.23867Mohm | 0 | 37.2474 | 109.83635kohm | 0.00346 | 0.93885 | 114963.00306Mohm | 0 | 35.2004 | 110.02964kohm | 0.00104 | 1e-05 | 12777.00009Mohm | 0 | 36.9544 | 110.78686kohm | 0.00346 | 1e-05 | 20911.18175Mohm | 0 | 35.0172 | 111.56619kohm | 0.94902 | 0 | 28744.87674Mohm | 0 | 36.4337 | 109.33661kohm | 0.00149 | 0.84463 | 19163.41799Mohm | 0 | 35.2819 | 110.04095kohm | 0.01031 | 0.84463 | 15333.19926Mohm | 0 | 35.7054 | 104.67518kohm | 0.00208 | 0.84463 | 19167.58466Mohm | 0 | 35.1931 | 110.74281kohm | 0.00208 | 0 | 13528.11768Mohm | 0 | 37.6023 | 110.59495kohm | 0.94902 | 1e-05 | 8212.28547Mohm | 0 | 35.6892 | 106.32103kohm | 0.0026 | 0.84463 | 11499.64949Mohm | 0 | 36.6579 | 114.09651kohm | 0.00149 | 23.8898 | 19168.08311Mohm | 0 | 35.0301 | 109.49702kohm | 0.00104 | 0.84463
0005_amshi_xlwritecheck | 08/03/2026 | 11:11:02 AM
lke | 08/19/2026 | 12:50:45 PM | 23.9918 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0 | 12 | 11.9917 | 0 | 0 | 15 | 14.9911 | 0 | 0 | 18 | 17.9909 | 0 | 0 | 21 | 20.9898 | 0 | 0 | 24 | 23.9915 | 0 | 0
lke | 08/19/2026 | 01:28:22 PM | 23.9917 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0 | 12 | 11.991 | 0 | 0 | 15 | 14.9905 | 0 | 0 | 18 | 17.9905 | 0 | 0 | 21 | 20.9896 | 0 | 0 | 24 | 23.9912 | 0 | 0
lke | 08/19/2026 | 02:00:41 PM | 23.9911 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 12 | 11.9897 | 0 | 0 | 15 | 14.9895 | 0 | 0 | 18 | 17.9898 | 0 | 0 | 21 | 20.989 | 0 | 0 | 24 | 23.991 | 0 | 0 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 |  |  |  |  | 18177.9997Mohm | 0 | 17.5987 | 105578.94588kohm | 0.95592 | 1 | 22215.44372Mohm | 0 | 15.7113 | 147769.22915kohm | 0.95592 | 1 | 16668.24988Mohm | 0 | 16.2138 | 100900.00002kohm | 0.95592 | 1 | 12499.87516Mohm | 0 | 15.6364 | 96095.24255kohm | 0.95592 | 1 | 13329.79928Mohm | 0 | 14.5895 | 89954.5501kohm | 0.95592 | 1 | 49987.50112Mohm | 0 | 15.753 | 197700.00163kohm | 0 | 1 | 13332.26591Mohm | 0 | 16.3162 | 88217.39368kohm | 0.95592 | 1 | 15386.38503Mohm | 0 | 15.6517 | 112944.43953kohm | 0.95592 | 1 | 66656.00328Mohm | 0 | 16.3189 | 179181.81083kohm | 0.95592 | 1 | 19996.70014Mohm | 0 | 14.1506 | 114764.70602kohm | 0.95592 | 62.0957 | 22222.33256Mohm | 0 | 15.9391 | 115294.11868kohm | 0.95592 | 0 | 16667.16753Mohm | 0 | 14.0531 | 113444.44036kohm | 0.95592 | 1 | 16664.50029Mohm | 0 | 16.3789 | 104684.20863kohm | 0.95592 | 1 | 6894.8623Mohm | 0 | 15.2394 | 54371.43146kohm | 0.95592 | 1 | 16665.58367Mohm | 0 | 16.4198 | 109842.11106kohm | 0.95592 | 1 | 18181.81828Mohm | 0 | 15.6136 | 124125.00659kohm | 0.95592 | 43.9101
lke | 08/20/2026 | 08:15:18 PM
lke | 08/20/2026 | 08:24:31 PM
lke | 08/20/2026 | 08:42:20 PM
lke | 08/20/2026 | 08:49:15 PM
lke | 08/24/2026 | 03:48:05 PM
lke | 08/24/2026 | 05:01:29 PM
lke | 08/24/2026 | 05:09:20 PM
lke | 08/25/2026 | 01:28:04 PM
lke | 08/25/2026 | 03:28:34 PM
lke | 08/25/2026 | 03:37:49 PM
lke | 08/25/2026 | 07:12:33 PM
demo0901 | 09/01/2026 | 11:11:39 AM | 23.9912 | 1.0888 | 262.863 | 263.5335 | 254.727 | 267.4709 | 253.4405 | 266.3777 | 12 | 11.9905 | 0.4938 | 150.3329 | 15 | 14.99 | 0.6273 | 183.1399 | 18 | 17.9901 | 0.7659 | 212.4873 | 21 | 20.9891 | 0.9125 | 239.2054 | 24 | 23.991 | 1.0707 | 263.7613 | 12 | 11.9905 | 0.4938 | 155.6419 | 15 | 14.99 | 0.6273 | 188.5532 | 18 | 17.9901 | 0.7659 | 216.1262 | 21 | 20.9891 | 0.9125 | 240.319 | 24 | 23.991 | 1.0707 | 262.8563 | 12 | 11.9905 | 0.4938 | 148.5301 | 15 | 14.99 | 0.6273 | 179.7916 | 18 | 17.9901 | 0.7659 | 207.5736 | 21 | 20.9891 | 0.9125 | 232.0383 | 24 | 23.991 | 1.0707 | 254.2186 | 12 | 11.9905 | 0.4938 | 153.039 | 15 | 14.99 | 0.6273 | 186.9175 | 18 | 17.9901 | 0.7659 | 216.8276 | 21 | 20.9891 | 0.9125 | 243.4689 | 24 | 23.991 | 1.0707 | 268.5004 | 12 | 11.9905 | 0.4938 | 148.9122 | 15 | 14.99 | 0.6273 | 178.9564 | 18 | 17.9901 | 0.7659 | 205.7471 | 21 | 20.9891 | 0.9125 | 230.5979 | 24 | 23.991 | 1.0707 | 253.2451 | 12 | 11.9905 | 0.4938 | 153.3456 | 15 | 14.99 | 0.6273 | 186.7 | 18 | 17.9901 | 0.7659 | 215.8596 | 21 | 20.9891 | 0.9125 | 242.2552 | 24 | 23.991 | 1.0707 | 266.2894
demo0901_1 | 09/01/2026 | 11:22:08 AM | [136 empty] | 9997.89963Mohm | 0 | 35.0628 | 111.90142kohm | 0.00445 | 0 | 11760.99985Mohm | 0 | 35.0619 | 109.35933kohm | 0 | 0.86602 | 9524.33356Mohm | 0 | 34.36 | 110.85702kohm | 0.95592 | 0 | 10526.78943Mohm | 0 | 34.8635 | 109.06923kohm | 0.00223 | 0.86602 | 11107.9443Mohm | 0 | 34.6365 | 110.7247kohm | 0.00056 | 0 | 22215.3325Mohm | 0 | 35.0161 | 111.46426kohm | 0.00056 | 0.95592 | 7999.75977Mohm | 0 | 35.0455 | 111.51908kohm | 0 | 0.95592 | 10527.68439Mohm | 0 | 35.0419 | 111.69722kohm | 0.00112 | 0 | 14283.3574Mohm | 0 | 34.5217 | 109.71173kohm | 0.00149 | 0.86602 | 12497.375Mohm | 0 | 36.072 | 108.4324kohm | 0.00056 | 0.86602 | 8000.04004Mohm | 0 | 34.8935 | 105.5402kohm | 0.00445 | 0.86602 | 8000.47998Mohm | 0 | 35.7764 | 109.54812kohm | 0.00112 | 0 | 8694.47802Mohm | 0 | 35.5425 | 111.00505kohm | 0.00223 | 0 | 5127.00026Mohm | 0 | 35.5959 | 106.6838kohm | 0.00223 | 0.86602 | 9090.13621Mohm | 0 | 33.9768 | 114.32777kohm | 0.00089 | 22.5402 | 13333.86584Mohm | 0 | 34.9253 | 109.58679kohm | 0.00223 | 0
demo0901_2 | 09/01/2026 | 11:53:35 AM
09082026v1 | 09/08/2026 | 02:51:58 PM | 23.9906 | 1.0848 | 263.4689 | 264.1469 | 252.2702 | 266.91 | 255.6964 | 267.0423 | 12 | 11.989 | 0.4913 | 150.5166 | 15 | 14.9891 | 0.6207 | 183.46 | 18 | 17.9892 | 0.7652 | 212.9013 | 21 | 20.9885 | 0.9123 | 239.7679 | 24 | 23.9906 | 1.0683 | 263.7766 | 12 | 11.989 | 0.4913 | 155.3725 | 15 | 14.9891 | 0.6207 | 188.2786 | 18 | 17.9892 | 0.7652 | 216.3147 | 21 | 20.9885 | 0.9123 | 239.5317 | 24 | 23.9906 | 1.0683 | 263.182 | 12 | 11.989 | 0.4913 | 148.129 | 15 | 14.9891 | 0.6207 | 178.6876 | 18 | 17.9892 | 0.7652 | 205.5241 | 21 | 20.9885 | 0.9123 | 229.9547 | 24 | 23.9906 | 1.0683 | 251.4533 | 12 | 11.989 | 0.4913 | 152.4072 | 15 | 14.9891 | 0.6207 | 186.2312 | 18 | 17.9892 | 0.7652 | 215.8672 | 21 | 20.9885 | 0.9123 | 242.3444 | 24 | 23.9906 | 1.0683 | 266.6781 | 12 | 11.989 | 0.4913 | 150.4291 | 15 | 14.9891 | 0.6207 | 181.6237 | 18 | 17.9892 | 0.7652 | 208.9948 | 21 | 20.9885 | 0.9123 | 233.7003 | 24 | 23.9906 | 1.0683 | 256.7435 | 12 | 11.989 | 0.4913 | 154.0004 | 15 | 14.9891 | 0.6207 | 187.3195 | 18 | 17.9892 | 0.7652 | 216.7398 | 21 | 20.9885 | 0.9123 | 243.1186 | 24 | 23.9906 | 1.0683 | 267.0989 | [8 empty] | 13330.99963Mohm | 0 | 35.9796 | 113.70977kohm | 0 | 1e-05 | 12496.37555Mohm | 0 | 34.1893 | 109.0575kohm | 0.00519 | 0.94902 | 11766.11747Mohm | 0 | 36.1769 | 110.34407kohm | 0.00056 | 0 | 10001.14975Mohm | 0 | 33.5813 | 108.38801kohm | 0.00346 | 0.94902 | 12496.18787Mohm | 0 | 36.3224 | 110.0465kohm | 0 | 1e-05 | 33323.00083Mohm | 0 | 33.7105 | 110.88158kohm | 0.00173 | 1e-05 | 10525.8951Mohm | 0 | 36.8513 | 111.1719kohm | 0.00149 | 1e-05 | 14288.14255Mohm | 0 | 34.2808 | 111.1595kohm | 0.00208 | 0 | 28566.85779Mohm | 0 | 36.372 | 109.08584kohm | 0.95592 | 0.94902 | 16663.50033Mohm | 0 | 34.6664 | 110.18964kohm | 0.00208 | 0.94902 | 14285.57123Mohm | 0 | 35.9867 | 104.82548kohm | 0.00445 | 0 | 14287.07095Mohm | 0 | 34.0082 | 109.55407kohm | 0.00174 | 0 | 14284.00001Mohm | 0 | 36.7387 | 111.47762kohm | 0.00445 | 1e-05 | 6664.79954Mohm | 0 | 34.5492 | 107.19473kohm | 0.00208 | 0.94902 | 12499.12519Mohm | 0 | 35.9055 | 114.44112kohm | 0.95592 | 1 | 12501.24997Mohm | 0 | 33.4918 | 109.17722kohm | 0.00346 | 0
20260908v1 | 09/08/2026 | 03:44:09 PM | 23.9905 | 1.0784 | 263.6837 | 263.4942 | 251.5303 | 266.9505 | 254.9952 | 267.6305 | 12 | 11.9888 | 0.4894 | 150.604 | 15 | 14.9889 | 0.6213 | 183.5137 | 18 | 17.9889 | 0.7622 | 213.1816 | 21 | 20.9884 | 0.9064 | 239.8243 | 24 | 23.9902 | 1.0618 | 263.86 | 12 | 11.9888 | 0.4894 | 155.6636 | 15 | 14.9889 | 0.6213 | 188.5934 | 18 | 17.9889 | 0.7622 | 217.2894 | 21 | 20.9884 | 0.9064 | 242.6975 | 24 | 23.9902 | 1.0618 | 266.6063 | 12 | 11.9888 | 0.4894 | 148.3623 | 15 | 14.9889 | 0.6213 | 178.9354 | 18 | 17.9889 | 0.7622 | 205.6419 | 21 | 20.9884 | 0.9064 | 229.8867 | 24 | 23.9902 | 1.0618 | 251.8596 | 12 | 11.9888 | 0.4894 | 153.0768 | 15 | 14.9889 | 0.6213 | 186.6102 | 18 | 17.9889 | 0.7622 | 216.318 | 21 | 20.9884 | 0.9064 | 243.2859 | 24 | 23.9902 | 1.0618 | 267.9932 | 12 | 11.9888 | 0.4894 | 150.4909 | 15 | 14.9889 | 0.6213 | 181.4111 | 18 | 17.9889 | 0.7622 | 208.2694 | 21 | 20.9884 | 0.9064 | 233.0353 | 24 | 23.9902 | 1.0618 | 255.9218 | 12 | 11.9888 | 0.4894 | 154.1871 | 15 | 14.9889 | 0.6213 | 186.9619 | 18 | 17.9889 | 0.7622 | 216.3873 | 21 | 20.9884 | 0.9064 | 243.5442 | 24 | 23.9902 | 1.0618 | 267.5245 | 9.9e+37 | 9.9e+37 | 9.9e+37 | 9.9e+37 |  |  |  |  | 11762.52932Mohm | 0 | 35.4625 | 111.79951kohm | 0.00149 | 0.95174 | 15379.76976Mohm | 0 | 35.4568 | 108.7356kohm | 0.0005 | 0.95548 | 10000.99961Mohm | 0 | 35.5378 | 110.79685kohm | 0.00089 | 0 | 10527.47366Mohm | 0 | 34.7925 | 109.43172kohm | 0.00064 | 0 | 11107.99923Mohm | 0 | 35.0112 | 109.69044kohm | 0.00149 | 0 | 28564.1426Mohm | 0 | 35.1035 | 109.8944kohm | 0.00089 | 0 | 11110.77701Mohm | 0 | 36.2126 | 110.7886kohm | 0.0005 | 0.95174 | 14287.92893Mohm | 0 | 35.5339 | 111.86035kohm | 0.00149 | 0 | 19996.70014Mohm | 0 | 34.9627 | 110.95592kohm | 0.00112 | 0.95548 | 16663.16667Mohm | 0 | 35.8491 | 110.42978kohm | 0.00074 | 0.95548 | 14285.71422Mohm | 0 | 35.4987 | 104.93925kohm | 0.00223 | 0.95548 | 14286.85733Mohm | 0 | 36.1489 | 111.08096kohm | 0.00089 | 0 | 11109.66617Mohm | 0 | 36.1189 | 110.53812kohm | 0.00074 | 0 | 6664.96637Mohm | 0 | 35.7712 | 108.93745kohm | 0.00149 | 0.95548 | 12499.56312Mohm | 0 | 35.2988 | 114.41324kohm | 0.00223 | 0 | 12500.93793Mohm | 0 | 34.7028 | 109.49263kohm | 0.0005 | 0.95548
lke | 09/10/2026 | 02:07:13 PM